## General
Units
ppb | µg/L
cps | counts per second
Analysis of sample run using standard-sample data files (raw data) from:
19th Nov 2019
08th Jan 2020
27th Feb 2020
Note:
Concentrations selected for the calibration leading to leverage or bias. High concentration STDs are the point of influence because the disproportionately affect the position of the regression line.
Suggested Changes (Kathrin)
1. Standards with at least seven different concentrations (including a blank)
2. Standard concentrations cover the range of sample concentrations and are evenly spaced across the calibration range to avoid leverage/bias
3. Lowest calibration standard should be approximately equal to estimated LOQ and the remaining standards cover the full range of sample concentrations

## cobalt_Co
Date | No_cal | Std_conc (ppb) | Co 59 (cps) | Co_59 (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 66 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 465 | 399 | Multiple R | 0.99998 | Regression | 1 | 35458641.7254 | 35458641.7254 | 94887.4612 | 7.5447e-08 | Intercept | -71.8397 | 22.6706 | 50.693 | -3.1688 | 0.050526 | -143.9877 | 0.30836 | -143.9877 | 0.30836 | 0.0035456 | 0.010744
2020-01-08 | 1 | 0.02 | 901 | 835 | R Square | 0.99997 | Residual | 3 | 1121.0746 | 373.6915 |  |  | X Variable 1 | 47181.5232 | 220.1965 |  | 214.2701 | 2.2416e-07 | 46480.7597 | 47882.2868 | 46480.7597 | 47882.2868
2020-01-08 | 1 | 0.05 | 2304.3787 | 2238.3787 | Adjusted R Square | 0.99996 | Total | 4 | 35459762.8
2020-01-08 | 1 | 0.1 | 4706.6079 | 4640.6079 | Standard Error | 19.3311
2020-01-08 | 1 | 0.2 | 9442.5744 | 9376.5744 | Observations | 5
2020-01-08 | 2 | 0 | 56 | 0 | Regression Statistics |  | Regression | 1 | 53693438.7401 | 53693438.7401 | 45911.6699 | 2.2416e-07 | Intercept | -7.7761 | 12.8151 | 28.6554 | -0.6068 | 0.58683 | -48.5594 | 33.0071 | -48.5594 | 33.0071 | 0.0024663 | 0.0074737
2020-01-08 | 2 | 0.01 | 412 | 356 | Multiple R | 0.99997 | Residual | 3 | 3508.4831 | 1169.4944 |  |  | X Variable 1 | 38341.791 | 124.4709 |  | 308.0381 | 7.5447e-08 | 37945.6689 | 38737.9131 | 37945.6689 | 38737.9131
2020-01-08 | 2 | 0.02 | 811 | 755 | R Square | 0.99993 | Total | 4 | 53696947.2232
2020-01-08 | 2 | 0.05 | 1988 | 1932 | Adjusted R Square | 0.99991
2020-01-08 | 2 | 0.1 | 3893 | 3837 | Standard Error | 34.1979
2020-01-08 | 2 | 0.2 | 7707 | 7651 | Observations | 5
2020-01-08 | 3 | 0 | 58 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 400 | 342 | Multiple R | 0.99998 | Regression | 1 | 32289401.7439 | 32289401.7439 | 89737.9796 | 8.2033e-08 | Intercept | -18.3051 | 12.5749 | 28.1184 | -1.4557 | 0.24149 | -58.3242 | 21.714 | -58.3242 | 21.714 | 0.0025361 | 0.0076851
2020-01-08 | 3 | 0.02 | 783 | 725 | R Square | 0.99997 | Residual | 3 | 1079.4561 | 359.8187 |  |  | X Variable 1 | 36588.2255 | 122.1387 |  | 299.563 | 8.2033e-08 | 36199.5258 | 36976.9253 | 36199.5258 | 36976.9253
2020-01-08 | 3 | 0.05 | 1880 | 1822 | Adjusted R Square | 0.99996 | Total | 4 | 32290481.2
2020-01-08 | 3 | 0.1 | 3672 | 3614 | Standard Error | 18.9689
2020-01-08 | 3 | 0.2 | 7367 | 7309 | Observations | 5
2019-11-20 | 4 | 0 | 84 | 0 | Regression Statistics |  | Regression | 1 | 19014121.9171 | 19014121.9171 | 4154.1357 | 8.2295e-06 | Intercept | 87.6235 | 44.85 | 100.2876 | 1.9537 | 0.14575 | -55.109 | 230.3561 | -55.109 | 230.3561 | 0.011787 | 0.035719
2019-11-20 | 4 | 0.01 | 403.7557 | 319.7557 | Multiple R | 0.99964 | Residual | 3 | 13731.4643 | 4577.1548 |  |  | X Variable 1 | 28076.9213 | 435.6213 |  | 64.4526 | 8.2295e-06 | 26690.5798 | 29463.2629 | 26690.5798 | 29463.2629
2019-11-20 | 4 | 0.02 | 697.5171 | 613.5171 | R Square | 0.99928 | Total | 4 | 19027853.3814
2019-11-20 | 4 | 0.05 | 1633.3434 | 1549.3434 | Adjusted R Square | 0.99904
2019-11-20 | 4 | 0.1 | 3049.0762 | 2965.0762 | Standard Error | 67.6547
2019-11-20 | 4 | 0.2 | 5743.6554 | 5659.6554 | Observations | 5
2020-02-27 | 5 | 0 | 78.5002 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 5 | 0.01 | 539.5102 | 461.01 | Multiple R | 1 | Regression | 1 | 56166781.0388 | 56166781.0388 | 567349.5739 | 5.1605e-09 | Intercept | -16.5246 | 6.596 | 14.749 | -2.5053 | 0.087302 | -37.5159 | 4.4667 | -37.5159 | 4.4667 | 0.0010086 | 0.0030564
2020-02-27 | 5 | 0.02 | 1021.5365 | 943.0363 | R Square | 0.99999 | Residual | 3 | 296.9956 | 98.9985 |  |  | X Variable 1 | 48255.9773 | 64.0657 |  | 753.2261 | 5.1605e-09 | 48052.0916 | 48459.863 | 48052.0916 | 48459.863
2020-02-27 | 5 | 0.05 | 2478.9651 | 2400.4649 | Adjusted R Square | 0.99999 | Total | 4 | 56167078.0344
2020-02-27 | 5 | 0.1 | 4900.8409 | 4822.3407 | Standard Error | 9.9498
2020-02-27 | 5 | 0.2 | 9706.2968 | 9627.7966 | Observations | 5
2020-02-27 | 6 | 0 | 81.7502 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 6 | 0.01 | 491.0084 | 409.2582 | Multiple R | 0.99994 | Regression | 1 | 41443234.283 | 41443234.283 | 25969.3404 | 5.2689e-07 | Intercept | 3.3574 | 26.4826 | 59.2169 | 0.12678 | 0.90714 | -80.9221 | 87.6368 | -80.9221 | 87.6368 | 0.0047143 | 0.014286
2020-02-27 | 6 | 0.02 | 891.2778 | 809.5276 | R Square | 0.99988 | Residual | 3 | 4787.5572 | 1595.8524 |  |  | X Variable 1 | 41451.3013 | 257.2218 |  | 161.1501 | 5.2689e-07 | 40632.7068 | 42269.8957 | 40632.7068 | 42269.8957
2020-02-27 | 6 | 0.05 | 2215.4218 | 2133.6716 | Adjusted R Square | 0.99985 | Total | 4 | 41448021.8402
2020-02-27 | 6 | 0.1 | 4201.1177 | 4119.3674 | Standard Error | 39.9481
2020-02-27 | 6 | 0.2 | 8378.2068 | 8296.4566 | Observations | 5 | [15 empty] | MEAN | 0.004343 | 0.013161
|  |  |  | Concentration _calibration methods | [5 empty] | Concentrations relative to P.Elmer instrument data
Sample ID | Raw int (cps) | Ga_int |  | Calb_ no "0" intercept | Calb_ "0" intercept | Weighted | Conc_trend | Calb_ P.Elmer | Conc_Uncertainty _P.Elmer | Calb_ no "0" intercept | Calb_ "0" intercept | Weighted | Conc_trend
7014562 | 414.756 | 89.45% |  | 0.49814 | 0.41796 | 0.5216 | 0.5768 | 0.46499 | 5.323% | 6.6543 | 11.2515 | 10.8527 | 19.3839
5018625 | 323.0037 | 92.22% |  | 0.37767 | 0.29875 | 0.50592 | 0.45398 | 0.33232 | 4.862% | 12.0078 | 11.2387 | 34.3136 | 26.7981
5018625 | 308.7534 | 91.59% |  | 0.3638 | 0.28414 | 0.50943 | 0.44063 | 0.31724 | 11.2% | 12.7968 | 11.6494 | 37.7267 | 28.0036
6014755 | 483.0082 | 86.97% |  | 0.59556 | 0.51402 | 0.53648 | 0.67646 | 0.5724 | 4.826% | 3.8898 | 11.3577 | -6.6953 | 15.3827
6016405 | 450.7572 | 87.74% |  | 0.55136 | 0.4701 | 0.53177 | 0.63155 | 0.52383 | 10.06% | 4.993 | 11.4295 | 1.4929 | 17.0558
Blank 3 | 50
7017480 | 684.2664 | 89.5% |  | 0.81715 | 0.74053 | 0.5213 | 0.89572 | 0.81714 | 2.955% | 0.00102 | 10.3451 | -56.7506 | 8.7727
7017480 | 672.5158 | 88.71% |  | 0.81042 | 0.73296 | 0.52597 | 0.8897 | 0.81007 | 3.536% | 0.044011 | 10.5191 | -54.0139 | 8.9505
7017979 | 302.0032 | 88.54% |  | 0.36823 | 0.28575 | 0.52697 | 0.44772 | 0.32182 | 8.103% | 12.6041 | 12.6241 | 38.9298 | 28.1199
7017979 | 286.2529 | 88.31% |  | 0.35028 | 0.26737 | 0.52835 | 0.42998 | 0.3019 | 2.202% | 13.8124 | 12.9132 | 42.8592 | 29.7871
3010112 | 138.0007 | 91.83% |  | 0.16566 | 0.08405 | 0.50807 | 0.24232 | 0.098513 | 13.83% | 40.5326 | 17.2083 | 80.6101 | 59.3456
3010112 | 135.0006 | 90.92% |  | 0.16383 | 0.08136 | 0.51319 | 0.24127 | 0.096264 | 19.96% | 41.2422 | 18.3179 | 81.2421 | 60.1004
3014819 | 274.2526 | 88.6% |  | 0.33476 | 0.25197 | 0.52659 | 0.4142 | 0.28496 | 8.388% | 14.8764 | 13.0936 | 45.8861 | 31.2013
3014819 | 279.5027 | 88.36% |  | 0.34198 | 0.25903 | 0.52804 | 0.42163 | 0.29253 | 9.445% | 14.4607 | 12.9318 | 44.6012 | 30.62
3019179 | 330.7538 | 88.51% |  | 0.4028 | 0.32066 | 0.52714 | 0.4823 | 0.35982 | 5.756% | 10.6701 | 12.2105 | 31.7416 | 25.396
3019179 | 321.7536 | 88.97% |  | 0.38998 | 0.30815 | 0.5244 | 0.46907 | 0.34579 | 7.154% | 11.3302 | 12.2146 | 34.0596 | 26.2818
5011752 | 352.0043 | 88.48% |  | 0.4284 | 0.34652 | 0.52732 | 0.50793 | 0.38821 | 4.746% | 9.3815 | 12.0313 | 26.3806 | 23.5704
5011752 | 352.7544 | 89.3% |  | 0.42535 | 0.34423 | 0.52247 | 0.50415 | 0.38466 | 6.55% | 9.566 | 11.7448 | 26.3763 | 23.7009
Dilution factor | 50
Below lowest STD (0.01 ppb)

## nickel_Ni
Date | No_cal | Std_conc (ppb) | Ni 60 (cps) | Ni_60 (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 41 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.05 | 491.7341 | 450.7341 | Multiple R | 0.99999 | Regression | 1 | 53479639.0742 | 53479639.0742 | 170656.1796 | 3.1281e-08 | Intercept | -21.6057 | 11.7354 | 26.2411 | -1.8411 | 0.16286 | -58.9529 | 15.7416 | -58.9529 | 15.7416 | 0.0091952 | 0.027864
2020-01-08 | 1 | 0.1 | 954.4854 | 913.4854 | R Square | 0.99998 | Residual | 3 | 940.1295 | 313.3765 |  |  | X Variable 1 | 9417.4989 | 22.7968 |  | 413.1055 | 3.1281e-08 | 9344.9492 | 9490.0486 | 9344.9492 | 9490.0486
2020-01-08 | 1 | 0.25 | 2394 | 2353 | Adjusted R Square | 0.99998 | Total | 4 | 53480579.2038
2020-01-08 | 1 | 0.5 | 4707 | 4666 | Standard Error | 17.7024
2020-01-08 | 1 | 1 | 9443 | 9402 | Observations | 5
2020-01-08 | 2 | 0 | 51 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.05 | 432 | 381 | Multiple R | 0.99983 | Regression | 1 | 43006542.5433 | 43006542.5433 | 8908.5608 | 2.6217e-06 | Intercept | -37.7662 | 46.0604 | 102.9942 | -0.81993 | 0.47234 | -184.3509 | 108.8186 | -184.3509 | 108.8186 | 0.040246 | 0.12196
2020-01-08 | 2 | 0.1 | 884 | 833 | R Square | 0.99966 | Residual | 3 | 14482.6567 | 4827.5522 |  |  | X Variable 1 | 8445.1741 | 89.4756 |  | 94.3852 | 2.6217e-06 | 8160.4227 | 8729.9256 | 8160.4227 | 8729.9256
2020-01-08 | 2 | 0.25 | 2166 | 2115 | Adjusted R Square | 0.99955 | Total | 4 | 43021025.2
2020-01-08 | 2 | 0.5 | 4133 | 4082 | Standard Error | 69.4806
2020-01-08 | 2 | 1 | 8497 | 8446 | Observations | 5
2020-01-08 | 3 | 0 | 47 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.05 | 424 | 377 | Multiple R | 1 | Regression | 1 | 38980010.4081 | 38980010.4081 | 396686.7528 | 8.8266e-09 | Intercept | -13.8441 | 6.5715 | 14.6942 | -2.1067 | 0.12576 | -34.7574 | 7.0692 | -34.7574 | 7.0692 | 0.0060311 | 0.018276
2020-01-08 | 3 | 0.1 | 837 | 790 | R Square | 0.99999 | Residual | 3 | 294.7919 | 98.264 |  |  | X Variable 1 | 8040.1161 | 12.7655 |  | 629.8307 | 8.8266e-09 | 7999.4905 | 8080.7417 | 7999.4905 | 8080.7417
2020-01-08 | 3 | 0.25 | 2053 | 2006 | Adjusted R Square | 0.99999 | Total | 4 | 38980305.2
2020-01-08 | 3 | 0.5 | 4060 | 4013 | Standard Error | 9.9128
2020-01-08 | 3 | 1 | 8068 | 8021 | Observations | 5
2019-11-20 | 1 | 0 | 43.7 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.05 | 379.8 | 336.1 | Multiple R | 0.99968 | Regression | 1 | 24359586.1114 | 24359586.1114 | 4666.7899 | 6.9121e-06 | Intercept | 70.5222 | 47.895 | 107.0965 | 1.4724 | 0.23731 | -81.901 | 222.9455 | -81.901 | 222.9455 | 0.055605 | 0.1685
2019-11-20 | 1 | 0.1 | 709.6 | 665.9 | R Square | 0.99936 | Residual | 3 | 15659.3206 | 5219.7735 |  |  | X Variable 1 | 6355.8889 | 93.0395 |  | 68.3139 | 6.9121e-06 | 6059.7958 | 6651.982 | 6059.7958 | 6651.982
2019-11-20 | 1 | 0.25 | 1772.9 | 1729.2 | Adjusted R Square | 0.99914 | Total | 4 | 24375245.432
2019-11-20 | 1 | 0.5 | 3359.2 | 3315.5 | Standard Error | 72.248
2019-11-20 | 1 | 1 | 6425.8 | 6382.1 | Observations | 5
2020-02-27 | 1 | 0 | 24.4367 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.05 | 563.4616 | 539.0249 | Multiple R | 0.99989 | Regression | 1 | 69504135.5986 | 69504135.5986 | 14035.2277 | 1.326e-06 | Intercept | -33.8994 | 46.6509 | 104.3145 | -0.72666 | 0.52 | -182.3633 | 114.5646 | -182.3633 | 114.5646 | 0.032064 | 0.097162
2020-02-27 | 1 | 0.1 | 1109.489 | 1085.0524 | R Square | 0.99979 | Residual | 3 | 14856.3608 | 4952.1203 |  |  | X Variable 1 | 10736.1029 | 90.6227 |  | 118.4704 | 1.326e-06 | 10447.7011 | 11024.5048 | 10447.7011 | 11024.5048
2020-02-27 | 1 | 0.25 | 2631.8365 | 2607.3998 | Adjusted R Square | 0.99972 | Total | 4 | 69518991.9594
2020-02-27 | 1 | 0.5 | 5272.4378 | 5248.0011 | Standard Error | 70.3713
2020-02-27 | 1 | 1 | 10774.0573 | 10749.6206 | Observations | 5
2020-02-27 | 2 | 0 | 33.9671 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.05 | 514.9451 | 480.978 | Multiple R | 0.99998 | Regression | 1 | 49949241.6515 | 49949241.6515 | 78185.0107 | 1.0087e-07 | Intercept | 31.7611 | 16.7559 | 37.4673 | 1.8955 | 0.15431 | -21.5635 | 85.0858 | -21.5635 | 85.0858 | 0.013585 | 0.041167
2020-02-27 | 2 | 0.1 | 954.2246 | 920.2575 | R Square | 0.99996 | Residual | 3 | 1916.5787 | 638.8596 |  |  | X Variable 1 | 9101.3495 | 32.5495 |  | 279.6158 | 1.0087e-07 | 8997.7626 | 9204.9365 | 8997.7626 | 9204.9365
2020-02-27 | 2 | 0.25 | 2378.0537 | 2344.0866 | Adjusted R Square | 0.99995 | Total | 4 | 49951158.2302
2020-02-27 | 2 | 0.5 | 4610.9865 | 4577.0194 | Standard Error | 25.2757
2020-02-27 | 2 | 1 | 9162.9954 | 9129.0283 | Observations | 5 | [15 empty] | MEAN | 0.026121 | 0.079154
|  |  |  | Concentration _calibration methods | [5 empty] | Concentrations relative to P.Elmer instrument data
Sample ID | Raw intensity | Ga_int |  | Calb_ no "0" intercept | Calb_ "0" intercept | Weighted | Conc_trend | Calb_ P.Elmer | Conc_Uncertainty _P.Elmer | Calb_ no "0" intercept | Calb_ "0" intercept | Weighted | Conc_trend
7014562 | 231.8513 | 89.45% |  | 0.024462 | 1.1365 |  | 1.467 | 1.1715 |  | -4689.0307 | 3.0769 | #DIV/0! | -20.1445
5018625 | 222.3041 | 92.22% |  | 1.1313 | 1.0472 |  | 1.3679 | 1.0739 |  | 5.0735 | 2.55 |  | -21.4914
5018625 | 163.3884 | 91.59% |  | 0.79729 | 0.71182 |  | 1.0355 | 0.73671 |  | 7.5987 | 3.4957 |  | -28.8567
6014755 | 499.2449 | 86.97% |  | 2.892 | 2.8067 |  | 3.1428 | 2.866 |  | 0.89964 | 2.1123 |  | -8.809
6016405 | 317.3871 | 87.74% |  | 1.7651 | 1.678 |  | 2.0137 | 1.7254 |  | 2.2481 | 2.8248 |  | -14.3187
Blank 3 | 40.0541
7017480 | 221.5608 | 89.5% |  | 1.1613 | 1.0746 |  | 1.4051 | 1.1084 |  | 4.5557 | 3.1418 |  | -21.1151
7017480 | 235.047 | 88.71% |  | 1.2525 | 1.1652 |  | 1.4984 | 1.2033 |  | 3.9258 | 3.2687 |  | -19.6956
7017979 | 163.6389 | 88.54% |  | 0.82623 | 0.73783 |  | 1.0727 | 0.77124 |  | 6.6552 | 4.5292 |  | -28.1002
7017979 | 158.9257 | 88.31% |  | 0.80003 | 0.71133 |  | 1.0471 | 0.74467 |  | 6.9199 | 4.6871 |  | -28.8833
3010112 | 145.0467 | 91.83% |  | 0.689 | 0.60352 |  | 0.9266 | 0.62698 |  | 9.0012 | 3.8877 |  | -32.335
3010112 | 132.0987 | 90.92% |  | 0.62027 | 0.53375 |  | 0.86026 | 0.55877 |  | 9.914 | 4.6884 |  | -35.0463
3014819 | 215.8046 | 88.6% |  | 1.1386 | 1.0509 |  | 1.3848 | 1.088 |  | 4.4387 | 3.5292 |  | -21.4318
3014819 | 214.7646 | 88.36% |  | 1.1354 | 1.0476 |  | 1.3824 | 1.0851 |  | 4.4306 | 3.587 |  | -21.502
3019179 | 240.6503 | 88.51% |  | 1.2889 | 1.2015 |  | 1.5354 | 1.2417 |  | 3.6602 | 3.3456 |  | -19.127
3019179 | 267.3317 | 88.97% |  | 1.4416 | 1.355 |  | 1.6868 | 1.3945 |  | 3.268 | 2.911 |  | -17.3303
5011752 | 283.3441 | 88.48% |  | 1.5458 | 1.459 |  | 1.7924 | 1.5005 |  | 2.9326 | 2.8434 |  | -16.2864
5011752 | 279.2341 | 89.3% |  | 1.5071 | 1.421 |  | 1.7514 | 1.4598 |  | 3.1404 | 2.7265 |  | -16.6517
6013040 | 276.2016 | 93.73% |  | 1.6546 | 1.4952 |  | 1.6652 | 1.5108 |  | 8.6908 | 1.0416 |  | -9.2701
6013040 | 297.2363 | 94.47% |  | 1.7738 | 1.6163 |  | 1.7843 | 1.6291 |  | 8.1589 | 0.793 |  | -8.6964
6022995 | 298.9889 | 92.52% |  | 1.8225 | 1.6617 |  | 1.8331 | 1.6789 |  | 7.8789 | 1.0365 |  | -8.4145
6022995 | 281.3574 | 94.99% |  | 1.6648 | 1.5077 |  | 1.6752 | 1.5184 |  | 8.7914 | 0.71305 |  | -9.3585
8022461 | 154.0302 | 94.37% |  | 0.87548 | 0.71374 |  | 0.88621 | 0.73501 |  | 16.0455 | 2.9802 |  | -17.0618
8022461 | 182.2787 | 95.24% |  | 1.0434 | 0.88395 |  | 1.054 | 0.90083 |  | 13.6658 | 1.9101 |  | -14.5322
5021103 | 188.4294 | 101.1% |  | 1.0187 | 0.86871 |  | 1.0287 | 0.87109 |  | 14.4929 | 0.2743 |  | -15.3199
5021103 | 160.1679 | 98.31% |  | 0.87737 | 0.72229 |  | 0.88766 | 0.73161 |  | 16.6139 | 1.2906 |  | -17.5802
6022863 | 711.3641 | 91.37% |  | 4.5224 | 4.3716 |  | 4.5324 | 4.3672 |  | 3.4336 | -0.10196 |  | -3.6461
6022863 | 648.5745 | 91.6% |  | 4.1045 | 3.9522 |  | 4.1145 | 3.9499 |  | 3.7652 | -0.057497 |  | -4.0008
7016034 | 81.5922 | 97.02% |  | 0.40862 | 0.24931 |  | 0.41919 | 0.26656 |  | 34.7662 | 6.9173 |  | -36.4112
7016034 | 55.4093 | 95.14% |  | 0.25342 | 0.090243 |  | 0.26425 | 0.11615 |  | 54.1661 | 28.7098 |  | -56.0443
Blank 7 | 33.2091
7013655 | 142.4513 | 94.47% |  | 0.8507 | 0.68903 |  | 0.85548 | 0.66074 |  | 22.33 | -4.1062 |  | -22.7645
7013655 | 142.53 | 90.65% |  | 0.88711 | 0.71862 |  | 1.1094 | 0.69938 |  | 21.1623 | -2.6781 |  | -36.9566
7016050 | 371.1444 | 94.45% |  | 2.2873 | 2.132 |  | 2.5001 | 2.0893 |  | 8.6562 | -2.0037 |  | -16.4332
7016050 | 359.0048 | 93.03% |  | 2.2447 | 2.0867 |  | 2.4609 | 2.0493 |  | 8.7052 | -1.7933 |  | -16.7234
3015106 | 523.717 | 90.75% |  | 3.3778 | 3.2207 |  | 3.599 | 3.1779 |  | 5.9189 | -1.3289 |  | -11.7021
3015106 | 502.265 | 90.82% |  | 3.2352 | 3.0775 |  | 3.4563 | 3.036 |  | 6.1557 | -1.3487 |  | -12.1593
3017494 | 427.5354 | 90.55% |  | 2.7551 | 2.5948 |  | 2.977 | 2.5583 |  | 7.143 | -1.4066 |  | -14.0646
3017494 | 440.0607 | 91.55% |  | 2.8064 | 2.6482 |  | 3.0259 | 2.6077 |  | 7.0794 | -1.5286 |  | -13.8193
3014967 | 261.6488 | 89.25% |  | 1.6928 | 1.5252 |  | 1.9183 | 1.5015 |  | 11.2974 | -1.5512 |  | -21.7243
3014967 | 234.8583 | 89.97% |  | 1.5026 | 1.3356 |  | 1.7264 | 1.313 |  | 12.6206 | -1.6924 |  | -23.9454
3016170 | 288.3277 | 90.26% |  | 1.8492 | 1.6842 |  | 2.0721 | 1.657 |  | 10.3939 | -1.6198 |  | -20.0336
3016170 | 271.0365 | 90.35% |  | 1.7337 | 1.5685 |  | 1.9564 | 1.5412 |  | 11.1056 | -1.7395 |  | -21.2251
Blank 9 | 44.1067
3021530 | 4.8013 | 0.01087% |  | -1699.9149 | -2254.4803 |  | -1698.5729 | 345.6266 |  | 120.332 | -115.3307 |  | -120.3481
3021530 | 133.1467 | 98.75% |  | 0.62174 | 0.5621 |  | 0.62189 | 0.5755 |  | 7.4379 | 2.3838 |  | -7.4595
4010272 | 1628.1227 | 90.78% |  | 10.9252 | 10.878 |  | 10.9254 | 10.9858 |  | -0.55473 | 0.99149 |  | 0.55362
4010272 | 64.2316 | 89.33% |  | 0.20721 | 0.14045 |  | 0.20737 | 0.1794 |  | 13.4221 | 27.7328 |  | -13.4897
5012473 | 353.2328 | 87.17% |  | 2.2756 | 2.2107 |  | 2.2757 | 2.2715 |  | 0.17895 | 2.7499 |  | -0.1859
5012473 | 343.9651 | 88.43% |  | 2.1781 | 2.114 |  | 2.1782 | 2.1707 |  | 0.33632 | 2.6836 |  | -0.34348
5019737 | 107.2426 | 93.9% |  | 0.48221 | 0.41918 |  | 0.48236 | 0.44668 |  | 7.3666 | 6.5609 |  | -7.396
5019737 | 109.4243 | 93.95% |  | 0.49639 | 0.43343 |  | 0.49654 | 0.45998 |  | 7.3349 | 6.1267 |  | -7.3634
5022428 | 152.9455 | 87.78% |  | 0.83981 | 0.77296 |  | 0.83998 | 0.82242 |  | 2.0714 | 6.3989 |  | -2.0904
5022428 | 139.3816 | 87.81% |  | 0.74343 | 0.67643 |  | 0.74359 | 0.72537 |  | 2.4285 | 7.2363 |  | -2.4499
6012213 | 232.2485 | 86.68% |  | 1.4198 | 1.3531 |  | 1.42 | 1.4096 |  | 0.71964 | 4.1762 |  | -0.73099
6012213 | 116.4346 | 88.2% |  | 0.57824 | 0.51125 |  | 0.5784 | 0.5576 |  | 3.5692 | 9.0656 |  | -3.5963
Blank 11 | 25.3128
6017142 | 210.7585 | 89.38% |  | 1.3582 | 1.2935 |  | 1.3584 | 1.2095 |  | 10.9513 | -6.494 |  | -10.9616
6017142 | 215.0429 | 89.56% |  | 1.3852 | 1.3206 |  | 1.3854 | 1.2356 |  | 10.797 | -6.4363 |  | -10.8072
6022103 | 1620.8698 | 85.86% |  | 11.6354 | 11.5856 |  | 11.6355 | 11.577 |  | 0.50169 | -0.07388 |  | -0.50277
6022103 | 1595.1037 | 87.51% |  | 11.2325 | 11.1833 |  | 11.2326 | 11.1775 |  | 0.48958 | -0.051892 |  | -0.49069
4016912 | 414.7208 | 87.2% |  | 2.848 | 2.7841 |  | 2.8481 | 2.7142 |  | 4.6958 | -2.5097 |  | -4.701
4016912 | 384.1772 | 87.42% |  | 2.6232 | 2.5591 |  | 2.6233 | 2.4879 |  | 5.1585 | -2.784 |  | -5.1642
4016955 | 324.1878 | 84.82% |  | 2.2635 | 2.1967 |  | 2.2637 | 2.1273 |  | 6.0163 | -3.1586 |  | -6.0231
4016955 | 334.3236 | 84.39% |  | 2.3498 | 2.2828 |  | 2.3499 | 2.2148 |  | 5.7458 | -2.9793 |  | -5.7524
4022742 | 51.7416 | 87.27% |  | 0.25707 | 0.18881 |  | 0.25724 | 0.10081 |  | 60.787 | -46.6092 |  | -60.8122
4022742 | 38.1029 | 87.13% |  | 0.16006 | 0.091518 |  | 0.16023 | 0.0026526 |  | 98.3427 | -97.1015 |  | -98.3444
3021041 | 25.9696 | 88.64% |  | 0.072139 | 0.0046194 |  | 0.072303 | -0.087335 |  | 221.0647 | -1990.6311 |  | -220.7909
3021041 | 13.8084 | 87.85% |  | -0.013367 | -0.08164 |  | -0.013202 | -0.17192 |  | -1186.1182 | 110.5822 |  | 1202.224
Dilutionfactor | 50
Below lowest STD (0.05 ppb)

## zinc_Zn
Date | No_cal | Std_conc (ppb) | Zn 64 (cps) | Zn_64 (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 542 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 1 | 5778 | 5236 | Multiple R | 0.99998 | Regression | 1 | 7093544304.4313 | 7093544304.4313 | 85727.1589 | 8.7857e-08 | Intercept | -358.7214 | 190.6939 | 426.4044 | -1.8811 | 0.15652 | -965.5944 | 248.1516 | -965.5944 | 248.1516 | 0.25947 | 0.78628
2020-01-08 | 1 | 2 | 11162 | 10620 | R Square | 0.99997 | Residual | 3 | 248236.7687 | 82745.5896 |  |  | X Variable 1 | 5423.0423 | 18.5218 |  | 292.792 | 8.7857e-08 | 5364.0976 | 5481.987 | 5364.0976 | 5481.987
2020-01-08 | 1 | 5 | 27177 | 26635 | Adjusted R Square | 0.99995 | Total | 4 | 7093792541.2
2020-01-08 | 1 | 10 | 54031 | 53489 | Standard Error | 287.6553
2020-01-08 | 1 | 20 | 108844 | 108302 | Observations | 5
2020-01-08 | 2 | 0 | 714 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 1 | 5265 | 4551 | Multiple R | 0.99995 | Regression | 1 | 5597999555.5881 | 5597999555.5881 | 32018.9648 | 3.8487e-07 | Intercept | -361.9104 | 277.1897 | 619.815 | -1.3056 | 0.28278 | -1244.0517 | 520.2308 | -1244.0517 | 520.2308 | 0.42457 | 1.2866
2020-01-08 | 2 | 2 | 10120 | 9406 | R Square | 0.99991 | Residual | 3 | 524501.6119 | 174833.8706 |  |  | X Variable 1 | 4817.5672 | 26.923 |  | 178.9384 | 3.8487e-07 | 4731.886 | 4903.2483 | 4731.886 | 4903.2483
2020-01-08 | 2 | 5 | 24585 | 23871 | Adjusted R Square | 0.99988 | Total | 4 | 5598524057.2
2020-01-08 | 2 | 10 | 47890 | 47176 | Standard Error | 418.1314
2020-01-08 | 2 | 20 | 96968 | 96254 | Observations | 5
2020-01-08 | 3 | 0 | 679 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 1 | 5019 | 4340 | Multiple R | 0.99997 | Regression | 1 | 4920183866.7 | 4920183866.7 | 59663.5425 | 1.5131e-07 | Intercept | -106 | 190.3708 | 425.6821 | -0.55681 | 0.6165 | -711.845 | 499.845 | -711.845 | 499.845 | 0.31103 | 0.9425
2020-01-08 | 3 | 2 | 9565 | 8886 | R Square | 0.99995 | Residual | 3 | 247396.5 | 82465.5 |  |  | X Variable 1 | 4516.5 | 18.4905 |  | 244.2612 | 1.5131e-07 | 4457.6551 | 4575.3449 | 4457.6551 | 4575.3449
2020-01-08 | 3 | 5 | 23512 | 22833 | Adjusted R Square | 0.99993 | Total | 4 | 4920431263.2
2020-01-08 | 3 | 10 | 45411 | 44732 | Standard Error | 287.1681
2020-01-08 | 3 | 20 | 90985 | 90306 | Observations | 5
2019-11-20 | 1 | 0 | 234.7519 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 1 | 8852.7423 | 8617.9904 | Multiple R | 0.9989 | Regression | 1 | 3612534127.1557 | 3612534127.1557 | 1360.2425 | 4.3843e-05 | Intercept | 2609.5955 | 1080.3441 | 2415.7228 | 2.4155 | 0.094548 | -828.5416 | 6047.7325 | -828.5416 | 6047.7325 | 2.0599 | 6.2421
2019-11-20 | 1 | 2 | 8852.7423 | 8617.9904 | R Square | 0.9978 | Residual | 3 | 7967404.8968 | 2655801.6323 |  |  | X Variable 1 | 3870.0567 | 104.9323 |  | 36.8815 | 4.3843e-05 | 3536.1153 | 4203.9981 | 3536.1153 | 4203.9981
2019-11-20 | 1 | 5 | 21627.3652 | 21392.6133 | Adjusted R Square | 0.99707 | Total | 4 | 3620501532.0525
2019-11-20 | 1 | 10 | 41450.0634 | 41215.3115 | Standard Error | 1629.663
2019-11-20 | 1 | 20 | 80500.9781 | 80266.2262 | Observations | 5
2020-02-27 | 1 | 0 | 345.0042 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 1 | 6003.0113 | 5658.0071 | Multiple R | 0.99998 | Regression | 1 | 8029443056.4531 | 8029443056.4531 | 67073.2907 | 1.2695e-07 | Intercept | -232.5082 | 229.3677 | 512.8817 | -1.0137 | 0.38538 | -962.4585 | 497.442 | -962.4585 | 497.442 | 0.29334 | 0.88892
2020-02-27 | 1 | 2 | 11681.0242 | 11336.02 | R Square | 0.99996 | Residual | 3 | 359134.4473 | 119711.4824 |  |  | X Variable 1 | 5769.711 | 22.2782 |  | 258.9851 | 1.2695e-07 | 5698.812 | 5840.61 | 5698.812 | 5840.61
2020-02-27 | 1 | 5 | 29121.4062 | 28776.402 | Adjusted R Square | 0.99994 | Total | 4 | 8029802190.9005
2020-02-27 | 1 | 10 | 57287.1534 | 56942.1492 | Standard Error | 345.9935
2020-02-27 | 1 | 20 | 115718.9024 | 115373.8982 | Observations | 5
2020-02-27 | 2 | 0 | 353.0044 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 1 | 5478.3008 | 5125.2964 | Multiple R | 0.99999 | Regression | 1 | 6043406705.8554 | 6043406705.8554 | 225822.4433 | 2.055e-08 | Intercept | 82.431 | 108.448 | 242.4972 | 0.7601 | 0.50247 | -262.6991 | 427.561 | -262.6991 | 427.561 | 0.15987 | 0.48446
2020-02-27 | 2 | 2 | 10379.0191 | 10026.0148 | R Square | 0.99999 | Residual | 3 | 80285.289 | 26761.763 |  |  | X Variable 1 | 5005.5552 | 10.5334 |  | 475.2078 | 2.055e-08 | 4972.0333 | 5039.0772 | 4972.0333 | 5039.0772
2020-02-27 | 2 | 5 | 25632.7295 | 25279.7252 | Adjusted R Square | 0.99998 | Total | 4 | 6043486991.1444
2020-02-27 | 2 | 10 | 50287.3696 | 49934.3652 | Standard Error | 163.5902
2020-02-27 | 2 | 20 | 100610.8571 | 100257.8528 | Observations | 5
|  |  |  | Concentration _calibration methods | [5 empty] | Concentrations relative to P.Elmer instrument data
Sample ID | Raw intensity | Ga_int |  | Calb_ no "0" intercept | Calb_ "0" intercept | Weighted | Conc_trend | Calb_ P.Elmer | Conc_Uncertainty _P.Elmer | Calb_ no "0" intercept | Calb_ "0" intercept | Weighted | Conc_trend
7014562 | 308637.7607 | 89.45% |  | 3182.7856 | 3190.7604 |  | 3182.6789 | 3223.4654 | 1.143% | -1.2781 | -1.025 |  | -1.2815
5018625 | 388599.7915 | 92.22% |  | 3887.7005 | 3898.0908 |  | 3887.5667 | 3938.1867 | 1.113% | -1.2986 | -1.0286 |  | -1.3021
5018625 | 383854.4601 | 91.59% |  | 3866.8414 | 3877.145 |  | 3866.7084 | 3917.4105 | 1.168% | -1.3078 | -1.0385 |  | -1.3112
6014755 | 98797.212 | 86.97% |  | 1045.7848 | 1046.5987 |  | 1045.7596 | 1057.0157 | 0.7482% | -1.0739 | -0.99532 |  | -1.0764
6016405 | 124700.7163 | 87.74% |  | 1309.2056 | 1310.9164 |  | 1309.1703 | 1323.9934 | 1.406% | -1.1295 | -0.99755 |  | -1.1323
Blank 3 | 594.5124
7017480 | 7406.6701 | 89.5% |  | 72.8467 | 70.5089 |  | 73.3997 | 71.097 | 2.721% | 2.4018 | -0.83408 |  | 3.1372
7017480 | 8147.3236 | 88.71% |  | 81.2083 | 78.8752 |  | 81.766 | 79.5944 | 1.856% | 1.9874 | -0.91179 |  | 2.6559
7017979 | 84572.1225 | 88.54% |  | 878.3466 | 878.6522 |  | 878.8751 | 887.7397 | 0.5647% | -1.0694 | -1.0343 |  | -1.0086
7017979 | 85643.4699 | 88.31% |  | 891.8501 | 892.1937 |  | 892.3796 | 901.4407 | 0.4043% | -1.0754 | -1.0364 |  | -1.0154
3010112 | 42234.3506 | 91.83% |  | 421.167 | 420.05 |  | 421.6927 | 424.1344 | 0.7408% | -0.70456 | -0.97237 |  | -0.57901
3010112 | 86270.995 | 90.92% |  | 872.6428 | 872.9977 |  | 873.1569 | 881.8442 | 1.157% | -1.0544 | -1.0134 |  | -0.99494
3014819 | 67389.5917 | 88.6% |  | 698.6697 | 698.3813 |  | 699.2047 | 705.5819 | 0.6365% | -0.98933 | -1.031 |  | -0.91206
3014819 | 72160.0537 | 88.36% |  | 750.4409 | 750.317 |  | 750.9755 | 758.0931 | 1.223% | -1.0197 | -1.0364 |  | -0.94779
3019179 | 127229.4075 | 88.51% |  | 1323.6277 | 1325.4093 |  | 1324.1395 | 1339.1338 | 0.712% | -1.1715 | -1.0355 |  | -1.1324
3019179 | 251309.0724 | 88.97% |  | 2604.4034 | 2610.4462 |  | 2604.8637 | 2637.5366 | 0.8151% | -1.2722 | -1.0378 |  | -1.2543
5011752 | 30577.443 | 88.48% |  | 315.4853 | 313.9225 |  | 316.0356 | 317.1216 | 1.515% | -0.51864 | -1.0191 |  | -0.34362
5011752 | 30764.3431 | 89.3% |  | 314.5114 | 312.9694 |  | 315.0565 | 316.1293 | 1.474% | -0.51443 | -1.0097 |  | -0.34052
Blank 5 | 506
6013040 | 88907.5699 | 93.73% |  | 983.2577 | 984.1524 |  | 983.1891 | 974.8717 | 0.7744% | 0.85287 | 0.94301 |  | 0.84596
6013040 | 90408.1911 | 94.47% |  | 992.1248 | 993.0745 |  | 992.055 | 983.6247 | 0.3723% | 0.85675 | 0.95157 |  | 0.84978
6022995 | 62609.6699 | 92.52% |  | 700.6798 | 700.4761 |  | 700.6403 | 693.5435 | 0.3169% | 1.0185 | 0.98971 |  | 1.0129
6022995 | 59067.1207 | 94.99% |  | 643.6221 | 643.2782 |  | 643.5876 | 636.7427 | 0.698% | 1.0689 | 1.016 |  | 1.0635
8022461 | 183868.2273 | 94.37% |  | 2022.8055 | 2027.6241 |  | 2022.6314 | 2009.7121 | 0.5306% | 0.64729 | 0.8834 |  | 0.63874
8022461 | 262596.125 | 95.24% |  | 2863.6707 | 2871.6735 |  | 2863.411 | 2846.7161 | 0.5013% | 0.59206 | 0.86909 |  | 0.58304
5021103 | 13813.9268 | 101.1% |  | 139.3905 | 137.3198 |  | 139.4052 | 134.5905 | 1.079% | 3.4435 | 1.9875 |  | 3.4537
5021103 | 12868.5456 | 98.31% |  | 133.385 | 131.2174 |  | 133.4011 | 128.6711 | 1.081% | 3.534 | 1.9406 |  | 3.5457
6022863 | 35434.148 | 91.37% |  | 400.2454 | 398.8777 |  | 400.2367 | 394.5144 | 1.042% | 1.4319 | 1.0939 |  | 1.4297
6022863 | 32986.0445 | 91.6% |  | 371.4579 | 369.9893 |  | 371.4521 | 365.8679 | 0.2651% | 1.5049 | 1.1139 |  | 1.5033
7016034 | 23755.2432 | 97.02% |  | 251.816 | 250.0578 |  | 251.8205 | 246.5952 | 0.8351% | 2.0733 | 1.3847 |  | 2.075
7016034 | 22419.83 | 95.14% |  | 242.1853 | 240.3376 |  | 242.1913 | 237.0715 | 0.3099% | 2.1115 | 1.359 |  | 2.1139
Blank 7 | 514
7013655 | 365715.9979 | 94.47% |  | 4021.4584 | 4286.6996 |  | 3576.6046 | 3999.575 | 0.7649% | 0.54417 | 6.698 |  | -11.826
7013655 | 376930.3794 | 90.65% |  | 4319.7784 | 4604.7901 |  | 3841.7641 | 4296.3 | 0.77% | 0.54351 | 6.6993 |  | -11.8314
7016050 | 67242.9079 | 94.45% |  | 737.2314 | 783.4458 |  | 659.7758 | 729.9539 | 0.4791% | 0.98715 | 6.8278 |  | -10.6367
7016050 | 64948.8484 | 93.03% |  | 722.8367 | 768.0461 |  | 647.0678 | 715.6796 | 0.1027% | 0.99015 | 6.8181 |  | -10.6035
3015106 | 28444.297 | 90.75% |  | 322.8296 | 341.2871 |  | 291.9363 | 317.5519 | 0.9769% | 1.6348 | 6.9546 |  | -8.7744
3015106 | 27075.3847 | 90.82% |  | 306.9207 | 324.3195 |  | 277.8032 | 301.7323 | 0.1293% | 1.6905 | 6.9645 |  | -8.6137
3017494 | 60560.1118 | 90.55% |  | 692.2315 | 735.3178 |  | 620.0256 | 685.2955 | 0.5254% | 1.002 | 6.8028 |  | -10.527
3017494 | 62164.9797 | 91.55% |  | 702.9541 | 746.7889 |  | 629.4921 | 695.9479 | 0.4547% | 0.99668 | 6.8079 |  | -10.5571
3014967 | 63466.6842 | 89.25% |  | 736.2094 | 782.1836 |  | 659.1604 | 729.1105 | 0.5133% | 0.96425 | 6.7852 |  | -10.612
3014967 | 61765.5059 | 89.97% |  | 710.6834 | 754.9801 |  | 636.4475 | 703.8005 | 0.9493% | 0.96849 | 6.7789 |  | -10.5826
3016170 | 53307.5324 | 90.26% |  | 610.9606 | 648.6169 |  | 547.8634 | 604.4307 | 1.313% | 1.0688 | 6.8124 |  | -10.3251
3016170 | 53714.0555 | 90.35% |  | 615.0041 | 652.9332 |  | 551.4491 | 608.4583 | 0.7977% | 1.0644 | 6.8116 |  | -10.3381
Blank 9 | 2425.206
3021530 | 14.75 | 0.01087% |  | -42834.0072 | -245930.8947 |  | -238002.5657 | 1563.4975
3021530 | 20033.0433 | 98.75% |  | 198.3145 | 197.7208 |  | 198.3207 | 219.236 | 2.206% | -10.5496 | -10.8816 |  | -10.5462
4010272 | 98796.9857 | 90.78% |  | 1176.7404 | 1177.2241 |  | 1176.7099 | 1208.3945 | 1.339% | -2.69 | -2.6478 |  | -2.6926
4010272 | 105346.8178 | 89.33% |  | 1277.0354 | 1277.6224 |  | 1277.0013 | 1309.7839 | 1.166% | -2.5644 | -2.5173 |  | -2.5672
5012473 | 53821.7088 | 87.17% |  | 653.9772 | 653.8092 |  | 653.9676 | 682.3103 | 1.235% | -4.3324 | -4.3592 |  | -4.334
5012473 | 54915.8572 | 88.43% |  | 658.4015 | 658.252 |  | 658.3915 | 686.2559 | 0.3401% | -4.2306 | -4.2543 |  | -4.2322
5019737 | 10424.3025 | 93.9% |  | 95.2262 | 94.4685 |  | 95.2371 | 116.6898 | 1.473% | -22.5396 | -23.5224 |  | -22.5255
5019737 | 10253.4288 | 93.95% |  | 93.1589 | 92.3992 |  | 93.1699 | 114.5895 | 0.8445% | -23.0043 | -24.0156 |  | -22.9898
5022428 | 25117.3128 | 87.78% |  | 287.2521 | 286.6596 |  | 287.2566 | 312.2777 | 1.016% | -8.7121 | -8.9368 |  | -8.7104
5022428 | 25572.8734 | 87.81% |  | 292.9132 | 292.3275 |  | 292.9174 | 317.9484 | 0.4356% | -8.547 | -8.7644 |  | -8.5454
6012213 | 70861.3926 | 86.68% |  | 875.3834 | 875.4704 |  | 875.3653 | 905.6949 | 0.704% | -3.4627 | -3.4524 |  | -3.4648
6012213 | 70126.2106 | 88.2% |  | 851.1531 | 851.2277 |  | 851.1356 | 880.6932 | 0.7541% | -3.4706 | -3.4615 |  | -3.4727
Blank 11 | 412
6017142 | 49687.7987 | 89.38% |  | 611.5428 | 586.3708 |  | 581.4822 | 613.5133 | 0.6872% | -0.32223 | -4.6289 |  | -5.5085
6017142 | 55280.775 | 89.56% |  | 679.4634 | 654.4241 |  | 649.4614 | 681.9736 | 0.7446% | -0.36943 | -4.2097 |  | -5.006
6022103 | 43062.3098 | 85.86% |  | 551.1704 | 524.8648 |  | 519.8793 | 552.8478 | 0.5128% | -0.30433 | -5.3315 |  | -6.3416
6022103 | 43413.6233 | 87.51% |  | 545.2136 | 519.4097 |  | 514.5129 | 546.7964 | 1.015% | -0.29031 | -5.2727 |  | -6.2746
4016912 | 55442.1347 | 87.2% |  | 699.9594 | 674.2427 |  | 669.1428 | 702.8463 | 0.7423% | -0.41243 | -4.2423 |  | -5.0368
4016912 | 56401.6504 | 87.42% |  | 710.3051 | 684.6693 |  | 679.5678 | 713.1769 | 0.5755% | -0.4043 | -4.1637 |  | -4.9457
4016955 | 58048.9827 | 84.82% |  | 753.6226 | 727.2251 |  | 721.9409 | 757.0231 | 0.5825% | -0.45121 | -4.0975 |  | -4.8594
4016955 | 55878.3341 | 84.39% |  | 728.963 | 702.3978 |  | 697.1214 | 732.221 | 0.8679% | -0.44693 | -4.2459 |  | -5.0349
4022742 | 106989.8998 | 87.27% |  | 1353.6785 | 1328.7512 |  | 1322.8609 | 1361.9849 | 0.8743% | -0.61361 | -2.5011 |  | -2.9575
4022742 | 120875.9904 | 87.13% |  | 1532.3684 | 1507.6088 |  | 1501.4947 | 1542.1584 | 0.9787% | -0.63888 | -2.2917 |  | -2.7082
3021041 | 27930.5298 | 88.64% |  | 344.7955 | 319.0926 |  | 314.493 | 344.618 | 0.3916% | 0.051472 | -7.9994 |  | -9.5789
3021041 | 27484.9165 | 87.85% |  | 342.2639 | 316.3244 |  | 311.6904 | 342.0882 | 0.5207% | 0.051336 | -8.1448 |  | -9.7526
Dilutionfactor | 50

## copper_Cu
Date | No_cal | Std_conc (ppb) | Cu65 (cps) | Cu65 (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 542 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.25 | 5778 | 5236 | Multiple R | 0.99998 | Regression | 1 | 7093544304.4313 | 7093544304.4313 | 85727.1589 | 8.7857e-08 | Intercept | -358.7214 | 190.6939 | 426.4044 | -1.8811 | 0.15652 | -965.5944 | 248.1516 | -965.5944 | 248.1516 | 0.064868 | 0.19657
2020-01-08 | 1 | 0.5 | 11162 | 10620 | R Square | 0.99997 | Residual | 3 | 248236.7687 | 82745.5896 |  |  | X Variable 1 | 21692.1692 | 74.0873 |  | 292.792 | 8.7857e-08 | 21456.3903 | 21927.948 | 21456.3903 | 21927.948
2020-01-08 | 1 | 1.25 | 27177 | 26635 | Adjusted R Square | 0.99995 | Total | 4 | 7093792541.2
2020-01-08 | 1 | 2.5 | 54031 | 53489 | Standard Error | 287.6553
2020-01-08 | 1 | 5 | 108844 | 108302 | Observations | 5
2020-01-08 | 2 | 0 | 714 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.25 | 5265 | 4551 | Multiple R | 0.99995 | Regression | 1 | 5597999555.5881 | 5597999555.5881 | 32018.9648 | 3.8487e-07 | Intercept | -361.9104 | 277.1897 | 619.815 | -1.3056 | 0.28278 | -1244.0517 | 520.2308 | -1244.0517 | 520.2308 | 0.10614 | 0.32164
2020-01-08 | 2 | 0.5 | 10120 | 9406 | R Square | 0.99991 | Residual | 3 | 524501.6119 | 174833.8706 |  |  | X Variable 1 | 19270.2687 | 107.6922 |  | 178.9384 | 3.8487e-07 | 18927.5441 | 19612.9932 | 18927.5441 | 19612.9932
2020-01-08 | 2 | 1.25 | 24585 | 23871 | Adjusted R Square | 0.99988 | Total | 4 | 5598524057.2
2020-01-08 | 2 | 2.5 | 47890 | 47176 | Standard Error | 418.1314
2020-01-08 | 2 | 5 | 96968 | 96254 | Observations | 5
2020-01-08 | 3 | 0 | 679 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.25 | 5019 | 4340 | Multiple R | 0.99997 | Regression | 1 | 4920183866.7 | 4920183866.7 | 59663.5425 | 1.5131e-07 | Intercept | -106 | 190.3708 | 425.6821 | -0.55681 | 0.6165 | -711.845 | 499.845 | -711.845 | 499.845 | 0.077757 | 0.23563
2020-01-08 | 3 | 0.5 | 9565 | 8886 | R Square | 0.99995 | Residual | 3 | 247396.5 | 82465.5 |  |  | X Variable 1 | 18066 | 73.9618 |  | 244.2612 | 1.5131e-07 | 17830.6205 | 18301.3795 | 17830.6205 | 18301.3795
2020-01-08 | 3 | 1.25 | 23512 | 22833 | Adjusted R Square | 0.99993 | Total | 4 | 4920431263.2
2020-01-08 | 3 | 2.5 | 45411 | 44732 | Standard Error | 287.1681
2020-01-08 | 3 | 5 | 90985 | 90306 | Observations | 5
2019-11-20 | 1 | 0 | 85.2503 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.25 | 2015.8922 | 1930.642 | Multiple R | 0.99989 | Regression | 1 | 778222218.1949 | 778222218.1949 | 13278.2269 | 1.4409e-06 | Intercept | 338.0379 | 160.4893 | 358.865 | 2.1063 | 0.12581 | -172.7106 | 848.7865 | -172.7106 | 848.7865 | 0.16482 | 0.49947
2019-11-20 | 1 | 0.5 | 3924.289 | 3839.0387 | R Square | 0.99977 | Residual | 3 | 175826.6884 | 58608.8961 |  |  | X Variable 1 | 7184.9401 | 62.3524 |  | 115.2312 | 1.4409e-06 | 6986.507 | 7383.3732 | 6986.507 | 7383.3732
2019-11-20 | 1 | 1.25 | 9607.9804 | 9522.7301 | Adjusted R Square | 0.9997 | Total | 4 | 778398044.8832
2019-11-20 | 1 | 2.5 | 18631.9065 | 18546.6562 | Standard Error | 242.0927
2019-11-20 | 1 | 5 | 36193.304 | 36108.0537 | Observations | 5
2020-02-27 | 1 | 0 | 113.0004 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.25 | 2734.2618 | 2621.2613 | Multiple R | 0.99996 | Regression | 1 | 1854738805.8832 | 1854738805.8832 | 41470.7987 | 2.6111e-07 | Intercept | -260.4205 | 140.1956 | 313.4869 | -1.8576 | 0.16022 | -706.5855 | 185.7446 | -706.5855 | 185.7446 | 0.093265 | 0.28262
2020-02-27 | 1 | 0.5 | 5449.54 | 5336.5396 | R Square | 0.99993 | Residual | 3 | 134171.9135 | 44723.9712 |  |  | X Variable 1 | 11092.073 | 54.468 |  | 203.6438 | 2.6111e-07 | 10918.7315 | 11265.4145 | 10918.7315 | 11265.4145
2020-02-27 | 1 | 1.25 | 13728.5954 | 13615.595 | Adjusted R Square | 0.9999 | Total | 4 | 1854872977.7967
2020-02-27 | 1 | 2.5 | 27268.2531 | 27155.2526 | Standard Error | 211.4804
2020-02-27 | 1 | 5 | 55456.9431 | 55343.9426 | Observations | 5
2020-02-27 | 2 | 0 | 161.2509 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.25 | 2444.9594 | 2283.7085 | Multiple R | 0.99999 | Regression | 1 | 1352561944.5211 | 1352561944.5211 | 247526.4054 | 1.7907e-08 | Intercept | -31.8005 | 49.0041 | 109.5764 | -0.64894 | 0.56263 | -187.7533 | 124.1523 | -187.7533 | 124.1523 | 0.038175 | 0.11568
2020-02-27 | 2 | 0.5 | 4824.3146 | 4663.0637 | R Square | 0.99999 | Residual | 3 | 16392.9413 | 5464.3138 |  |  | X Variable 1 | 9472.1795 | 19.0388 |  | 497.5203 | 1.7907e-08 | 9411.5896 | 9532.7694 | 9411.5896 | 9532.7694
2020-02-27 | 2 | 1.25 | 12037.8212 | 11876.5703 | Adjusted R Square | 0.99998 | Total | 4 | 1352578337.4624
2020-02-27 | 2 | 2.5 | 23881.7043 | 23720.4534 | Standard Error | 73.921
2020-02-27 | 2 | 5 | 47444.1585 | 47282.9076 | Observations | 5

## strontium_Sr
Date | No_cal | Std_conc (ppb) | Sr87 (cps) | Sr87 (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 139.7507 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.25 | 14178.2857 | 14038.535 | Multiple R | 0.99998 | Regression | 1 | 53798323928.0934 | 53798323928.0934 | 68467.3294 | 1.2309e-07 | Intercept | -1473.226 | 587.6341 | 1313.9897 | -2.507 | 0.087164 | -3343.3399 | 396.8878 | -3343.3399 | 396.8878 | 0.072586 | 0.21996
2020-01-08 | 1 | 0.5 | 29070.5548 | 28930.8041 | R Square | 0.99996 | Residual | 3 | 2357255.2512 | 785751.7504 |  |  | X Variable 1 | 59738.6924 | 228.3043 |  | 261.6626 | 1.2309e-07 | 59012.1264 | 60465.2585 | 59012.1264 | 60465.2585
2020-01-08 | 1 | 1.25 | 72576.9529 | 72437.2022 | Adjusted R Square | 0.99994 | Total | 4 | 53800681183.3446
2020-01-08 | 1 | 2.5 | 147098.7635 | 146959.0128 | Standard Error | 886.4264
2020-01-08 | 1 | 5 | 297925.6444 | 297785.8937 | Observations | 5
2020-01-08 | 2 | 0 | 186.0012 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.25 | 12956.1227 | 12770.1215 | Multiple R | 0.99999 | Regression | 1 | 41650607059.6163 | 41650607059.6163 | 131737.0783 | 4.6121e-08 | Intercept | -472.5253 | 372.7528 | 833.5006 | -1.2677 | 0.29437 | -1658.7911 | 713.7406 | -1658.7911 | 713.7406 | 0.052328 | 0.15857
2020-01-08 | 2 | 0.5 | 25817.3137 | 25631.3125 | R Square | 0.99998 | Residual | 3 | 948493.9457 | 316164.6486 |  |  | X Variable 1 | 52563.2255 | 144.8198 |  | 362.956 | 4.6121e-08 | 52102.3442 | 53024.1068 | 52102.3442 | 53024.1068
2020-01-08 | 2 | 1.25 | 65986.367 | 65800.3658 | Adjusted R Square | 0.99997 | Total | 4 | 41651555553.562
2020-01-08 | 2 | 2.5 | 130394.7491 | 130208.7479 | Standard Error | 562.2852
2020-01-08 | 2 | 5 | 262763.4694 | 262577.4681 | Observations | 5
2020-01-08 | 3 | 0 | 255.7523 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.25 | 12306.2988 | 12050.5465 | Multiple R | 0.99999 | Regression | 1 | 37143855602.5081 | 37143855602.5081 | 141332.3283 | 4.1505e-08 | Intercept | 34.0491 | 339.8498 | 759.9272 | 0.10019 | 0.92651 | -1047.5046 | 1115.6029 | -1047.5046 | 1115.6029 | 0.050521 | 0.15309
2020-01-08 | 3 | 0.5 | 24860.8643 | 24605.112 | R Square | 0.99998 | Residual | 3 | 788436.5037 | 262812.1679 |  |  | X Variable 1 | 49638.0636 | 132.0365 |  | 375.9419 | 4.1505e-08 | 49217.8644 | 50058.2627 | 49217.8644 | 50058.2627
2020-01-08 | 3 | 1.25 | 63062.134 | 62806.3817 | Adjusted R Square | 0.99997 | Total | 4 | 37144644039.0118
2020-01-08 | 3 | 2.5 | 124490.9099 | 124235.1576 | Standard Error | 512.6521
2020-01-08 | 3 | 5 | 248290.404 | 248034.6517 | Observations | 5
2019-11-20 | 1 | 0 | 95.7503 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.25 | 12118.1396 | 12022.3892 | Multiple R | 0.99978 | Regression | 1 | 28281459885.782 | 28281459885.782 | 6772.3079 | 3.9549e-06 | Intercept | 2851.0626 | 1354.7116 | 3029.2272 | 2.1046 | 0.12602 | -1460.2343 | 7162.3595 | -1460.2343 | 7162.3595 | 0.23079 | 0.69937
2019-11-20 | 1 | 0.5 | 23733.4526 | 23637.7023 | R Square | 0.99956 | Residual | 3 | 12528133.8885 | 4176044.6295 |  |  | X Variable 1 | 43313.3975 | 526.3249 |  | 82.294 | 3.9549e-06 | 41638.3969 | 44988.3982 | 41638.3969 | 44988.3982
2019-11-20 | 1 | 1.25 | 58870.094 | 58774.3437 | Adjusted R Square | 0.99941 | Total | 4 | 28293988019.6705
2019-11-20 | 1 | 2.5 | 113272.4551 | 113176.7047 | Standard Error | 2043.5373
2019-11-20 | 1 | 5 | 218217.2 | 218121.4497 | Observations | 5
2020-02-27 | 1 | 0 | 114.7505 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.25 | 16442.4595 | 16327.7091 | Multiple R | 1 | Regression | 1 | 67318275263.3513 | 67318275263.3513 | 702663.3908 | 3.7441e-09 | Intercept | -434.0219 | 205.1904 | 458.8198 | -2.1152 | 0.12475 | -1087.0295 | 218.9856 | -1087.0295 | 218.9856 | 0.022658 | 0.06866
2020-02-27 | 1 | 0.5 | 33212.0684 | 33097.318 | R Square | 1 | Residual | 3 | 287413.3311 | 95804.4437 |  |  | X Variable 1 | 66824.825 | 79.7194 |  | 838.2502 | 3.7441e-09 | 66571.1222 | 67078.5278 | 66571.1222 | 67078.5278
2020-02-27 | 1 | 1.25 | 82795.4855 | 82680.735 | Adjusted R Square | 0.99999 | Total | 4 | 67318562676.6824
2020-02-27 | 1 | 2.5 | 167047.6359 | 166932.8855 | Standard Error | 309.5229
2020-02-27 | 1 | 5 | 333741.8304 | 333627.08 | Observations | 5
2020-02-27 | 2 | 0 | 191.2513 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.25 | 15027.9005 | 14836.6492 | Multiple R | 0.99999 | Regression | 1 | 52327703269.5749 | 52327703269.5749 | 164915.1865 | 3.2928e-08 | Intercept | 260.0526 | 373.4219 | 834.9968 | 0.6964 | 0.53628 | -928.3426 | 1448.4478 | -928.3426 | 1448.4478 | 0.046769 | 0.14173
2020-02-27 | 2 | 0.5 | 29763.7264 | 29572.4752 | R Square | 0.99998 | Residual | 3 | 951902.0848 | 317300.6949 |  |  | X Variable 1 | 58916.5322 | 145.0798 |  | 406.0975 | 3.2928e-08 | 58454.8236 | 59378.2407 | 58454.8236 | 59378.2407
2020-02-27 | 2 | 1.25 | 74847.6487 | 74656.3974 | Adjusted R Square | 0.99998 | Total | 4 | 52328655171.6597
2020-02-27 | 2 | 2.5 | 147168.9524 | 146977.7012 | Standard Error | 563.2945
2020-02-27 | 2 | 5 | 295155.3469 | 294964.0956 | Observations | 5

## molybdenum_Mo
Date | No_cal | Std_conc (ppb) | Mo (cps) | Mo (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 6 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.5 | 5378 | 5372 | Multiple R | 0.99999 | Regression | 1 | 7750021950.2702 | 7750021950.2702 | 136393.9974 | 4.3779e-08 | Intercept | -390.8706 | 158.0222 | 353.3484 | -2.4735 | 0.089784 | -893.7678 | 112.0265 | -893.7678 | 112.0265 | 0.10285 | 0.31168
2020-01-08 | 1 | 1 | 11055 | 11049 | R Square | 0.99998 | Residual | 3 | 170462.5299 | 56820.8433 |  |  | X Variable 1 | 11336.8607 | 30.6969 |  | 369.3156 | 4.3779e-08 | 11239.1693 | 11434.5521 | 11239.1693 | 11434.5521
2020-01-08 | 1 | 2.5 | 27949 | 27943 | Adjusted R Square | 0.99997 | Total | 4 | 7750192412.8
2020-01-08 | 1 | 5 | 55947 | 55941 | Standard Error | 238.3712
2020-01-08 | 1 | 10 | 113147 | 113141 | Observations | 5
2020-01-08 | 2 | 0 | 19 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.5 | 4973.5609 | 4954.5609 | Multiple R | 0.99993 | Regression | 1 | 6537071778.432 | 6537071778.432 | 20786.7014 | 7.3573e-07 | Intercept | -365.8942 | 371.7604 | 831.2815 | -0.98422 | 0.39758 | -1549.0017 | 817.2133 | -1549.0017 | 817.2133 | 0.26347 | 0.79839
2020-01-08 | 2 | 1 | 10181.7958 | 10162.7958 | R Square | 0.99986 | Residual | 3 | 943450.091 | 314483.3637 |  |  | X Variable 1 | 10411.9715 | 72.2171 |  | 144.1759 | 7.3573e-07 | 10182.1444 | 10641.7986 | 10182.1444 | 10641.7986
2020-01-08 | 2 | 2.5 | 25964.9961 | 25945.9961 | Adjusted R Square | 0.99981 | Total | 4 | 6538015228.5231
2020-01-08 | 2 | 5 | 50862.4356 | 50843.4356 | Standard Error | 560.7882
2020-01-08 | 2 | 10 | 104110.1989 | 104091.1989 | Observations | 5
2020-01-08 | 3 | 0 | 52.6405 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.5 | 4921.766 | 4869.1255 | Multiple R | 0.99999 | Regression | 1 | 6080842376.4134 | 6080842376.4134 | 148617.8672 | 3.8491e-08 | Intercept | -116.1977 | 134.0944 | 299.8442 | -0.86654 | 0.44994 | -542.946 | 310.5505 | -542.946 | 310.5505 | 0.098534 | 0.29859
2020-01-08 | 3 | 1 | 9953.3296 | 9900.6891 | R Square | 0.99998 | Residual | 3 | 122747.8732 | 40915.9577 |  |  | X Variable 1 | 10042.0694 | 26.0488 |  | 385.5099 | 3.8491e-08 | 9959.1705 | 10124.9683 | 9959.1705 | 10124.9683
2020-01-08 | 3 | 2.5 | 25279.848 | 25227.2075 | Adjusted R Square | 0.99997 | Total | 4 | 6080965124.2866
2020-01-08 | 3 | 5 | 49902.622 | 49849.9816 | Standard Error | 202.2769
2020-01-08 | 3 | 10 | 100423.966 | 100371.3255 | Observations | 5
2019-11-20 | 1 | 0 | 11.2361 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.5 | 6170.4015 | 6159.1654 | Multiple R | 0.9999 | Regression | 1 | 7476136983.9211 | 7476136983.9211 | 14372.5398 | 1.2796e-06 | Intercept | 1215.4771 | 478.1192 | 1069.1069 | 2.5422 | 0.084516 | -306.1115 | 2737.0656 | -306.1115 | 2737.0656 | 0.31685 | 0.96015
2019-11-20 | 1 | 1 | 12066.5846 | 12055.3485 | R Square | 0.99979 | Residual | 3 | 1560504.3517 | 520168.1172 |  |  | X Variable 1 | 11134.7371 | 92.8781 |  | 119.8855 | 1.2796e-06 | 10839.1577 | 11430.3166 | 10839.1577 | 11430.3166
2019-11-20 | 1 | 2.5 | 29821.2901 | 29810.054 | Adjusted R Square | 0.99972 | Total | 4 | 7477697488.2728
2019-11-20 | 1 | 5 | 57479.5484 | 57468.3123 | Standard Error | 721.2268
2019-11-20 | 1 | 10 | 112155.7469 | 112144.5108 | Observations | 5
2020-02-27 | 1 | 0 | 12.9793 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.5 | 7520.959 | 7507.9796 | Multiple R | 1 | Regression | 1 | 13696622280.51 | 13696622280.51 | 483327.0945 | 6.5631e-09 | Intercept | 64.7602 | 111.5965 | 249.5373 | 0.58031 | 0.60242 | -290.3896 | 419.9099 | -290.3896 | 419.9099 | 0.054639 | 0.16557
2020-02-27 | 1 | 1 | 15259.8263 | 15246.8469 | R Square | 0.99999 | Residual | 3 | 85014.615 | 28338.205 |  |  | X Variable 1 | 15071.2087 | 21.6784 |  | 695.2173 | 6.5631e-09 | 15002.2183 | 15140.1991 | 15002.2183 | 15140.1991
2020-02-27 | 1 | 2.5 | 37879.8936 | 37866.9142 | Adjusted R Square | 0.99999 | Total | 4 | 13696707295.125
2020-02-27 | 1 | 5 | 75223.5352 | 75210.5559 | Standard Error | 168.3396
2020-02-27 | 1 | 10 | 150857.4486 | 150844.4693 | Observations | 5
2020-02-27 | 2 | 0 | 36.202 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.5 | 6942.8616 | 6906.6596 | Multiple R | 0.99995 | Regression | 1 | 10988847351.9089 | 10988847351.9089 | 28237.465 | 4.647e-07 | Intercept | 658.7531 | 413.5495 | 924.7248 | 1.5929 | 0.20943 | -657.3459 | 1974.8522 | -657.3459 | 1974.8522 | 0.22605 | 0.68501
2020-02-27 | 2 | 1 | 13904.4123 | 13868.2103 | R Square | 0.99989 | Residual | 3 | 1167475.2711 | 389158.4237 |  |  | X Variable 1 | 13499.4916 | 80.335 |  | 168.0401 | 4.647e-07 | 13243.83 | 13755.1533 | 13243.83 | 13755.1533
2020-02-27 | 2 | 2.5 | 35051.5755 | 35015.3735 | Adjusted R Square | 0.99986 | Total | 4 | 10990014827.18
2020-02-27 | 2 | 5 | 68756.1333 | 68719.9313 | Standard Error | 623.8256
2020-02-27 | 2 | 10 | 135310.1341 | 135273.9321 | Observations | 5

## antimony_Sb
Date | No_cal | Std_conc (ppb) | Sb (cps) | Sb (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 26.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 152.0008 | 125.5008 | Multiple R | 0.99991 | Regression | 1 | 3690764.8512 | 3690764.8512 | 17339.3762 | 9.6567e-07 | Intercept | -2.6113 | 9.6718 | 21.6267 | -0.26999 | 0.80468 | -33.3912 | 28.1686 | -33.3912 | 28.1686 | 0.0057695 | 0.017483
2020-01-08 | 1 | 0.02 | 252.7522 | 226.2522 | R Square | 0.99983 | Residual | 3 | 638.5636 | 212.8545 |  |  | X Variable 1 | 12369.9953 | 93.9405 |  | 131.6791 | 9.6567e-07 | 12071.0348 | 12668.9559 | 12071.0348 | 12668.9559
2020-01-08 | 1 | 0.05 | 658.7652 | 632.2652 | Adjusted R Square | 0.99977 | Total | 4 | 3691403.4148
2020-01-08 | 1 | 0.1 | 1261.3058 | 1234.8058 | Standard Error | 14.5895
2020-01-08 | 1 | 0.2 | 2495.2179 | 2468.7179 | Observations | 5
2020-01-08 | 2 | 0 | 32.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.01 | 147.5008 | 115.2507 | Multiple R | 0.99994 | Regression | 1 | 3140858.652 | 3140858.652 | 26083.4741 | 5.2343e-07 | Intercept | -3.3596 | 7.2745 | 16.2664 | -0.46182 | 0.67564 | -26.5104 | 19.7913 | -26.5104 | 19.7913 | 0.004704 | 0.014255
2020-01-08 | 2 | 0.02 | 244.2521 | 212.0021 | R Square | 0.99988 | Residual | 3 | 361.247 | 120.4157 |  |  | X Variable 1 | 11411.3111 | 70.6566 |  | 161.5038 | 5.2343e-07 | 11186.4502 | 11636.172 | 11186.4502 | 11636.172
2020-01-08 | 2 | 0.05 | 612.0131 | 579.7631 | Adjusted R Square | 0.99985 | Total | 4 | 3141219.8989
2020-01-08 | 2 | 0.1 | 1165.7976 | 1133.5475 | Standard Error | 10.9734
2020-01-08 | 2 | 0.2 | 2311.187 | 2278.937 | Observations | 5
2020-01-08 | 3 | 0 | 27.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 144.0007 | 116.7507 | Multiple R | 0.99981 | Regression | 1 | 3145922.4886 | 3145922.4886 | 7907.6539 | 3.1347e-06 | Intercept | -5.2087 | 13.2225 | 29.5665 | -0.39393 | 0.71996 | -47.2887 | 36.8713 | -47.2887 | 36.8713 | 0.0085433 | 0.025889
2020-01-08 | 3 | 0.02 | 224.0018 | 196.7517 | R Square | 0.99962 | Residual | 3 | 1193.4978 | 397.8326 |  |  | X Variable 1 | 11420.5063 | 128.4285 |  | 88.925 | 3.1347e-06 | 11011.7894 | 11829.2232 | 11011.7894 | 11829.2232
2020-01-08 | 3 | 0.05 | 605.0128 | 577.7628 | Adjusted R Square | 0.99949 | Total | 4 | 3147115.9863
2020-01-08 | 3 | 0.1 | 1179.0487 | 1151.7987 | Standard Error | 19.9457
2020-01-08 | 3 | 0.2 | 2297.9349 | 2270.6848 | Observations | 5
2019-11-20 | 1 | 0 | 41.5001 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.01 | 205.0015 | 163.5014 | Multiple R | 0.9999 | Regression | 1 | 7105412.5451 | 7105412.5451 | 14872.4085 | 1.2156e-06 | Intercept | -0.063256 | 14.49 | 32.4006 | -0.0043655 | 0.99679 | -46.1769 | 46.0504 | -46.1769 | 46.0504 | 0.0062296 | 0.018878
2019-11-20 | 1 | 0.02 | 369.5048 | 328.0047 | R Square | 0.9998 | Residual | 3 | 1433.2741 | 477.758 |  |  | X Variable 1 | 17163.5057 | 140.7393 |  | 121.9525 | 1.2156e-06 | 16715.6104 | 17611.4009 | 16715.6104 | 17611.4009
2019-11-20 | 1 | 0.05 | 932.5305 | 891.0304 | Adjusted R Square | 0.99973 | Total | 4 | 7106845.8192
2019-11-20 | 1 | 0.1 | 1751.1074 | 1709.6073 | Standard Error | 21.8577
2019-11-20 | 1 | 0.2 | 3471.1721 | 3429.672 | Observations | 5
2020-02-27 | 1 | 0 | 20.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.01 | 246.7521 | 226.5021 | Multiple R | 0.99999 | Regression | 1 | 11248999.3867 | 11248999.3867 | 124247.4549 | 5.0353e-08 | Intercept | 16.6503 | 6.3078 | 14.1047 | 2.6396 | 0.07768 | -3.4239 | 36.7245 | -3.4239 | 36.7245 | 0.0021553 | 0.0065312
2020-02-27 | 1 | 0.02 | 465.7576 | 445.5076 | R Square | 0.99998 | Residual | 3 | 271.6112 | 90.5371 |  |  | X Variable 1 | 21595.7502 | 61.2667 |  | 352.4875 | 5.0353e-08 | 21400.7722 | 21790.7282 | 21400.7722 | 21790.7282
2020-02-27 | 1 | 0.05 | 1119.2938 | 1099.0438 | Adjusted R Square | 0.99997 | Total | 4 | 11249270.9979
2020-02-27 | 1 | 0.1 | 2209.6709 | 2189.4209 | Standard Error | 9.5151
2020-02-27 | 1 | 0.2 | 4349.4121 | 4329.1621 | Observations | 5
2020-02-27 | 2 | 0 | 14 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.01 | 238.252 | 224.252 | Multiple R | 0.99996 | Regression | 1 | 8928184.0045 | 8928184.0045 | 38965.36 | 2.8669e-07 | Intercept | 14.2412 | 10.0347 | 22.4384 | 1.4192 | 0.2509 | -17.6939 | 46.1762 | -17.6939 | 46.1762 | 0.0038487 | 0.011663
2020-02-27 | 2 | 0.02 | 409.7559 | 395.7559 | R Square | 0.99992 | Residual | 3 | 687.3939 | 229.1313 |  |  | X Variable 1 | 19239.4616 | 97.4661 |  | 197.3965 | 2.8669e-07 | 18929.281 | 19549.6422 | 18929.281 | 19549.6422
2020-02-27 | 2 | 0.05 | 981.5337 | 967.5337 | Adjusted R Square | 0.9999 | Total | 4 | 8928871.3984
2020-02-27 | 2 | 0.1 | 1937.6315 | 1923.6314 | Standard Error | 15.1371
2020-02-27 | 2 | 0.2 | 3885.0282 | 3871.0282 | Observations | 5

## cesium_Cs
Date | No_cal | Std_conc (ppb) | Cs (cps) | Cs (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 3 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 453 | 450 | Multiple R | 0.99998 | Regression | 1 | 55203685.8932 | 55203685.8932 | 65998.7283 | 1.3006e-07 | Intercept | -33.2753 | 19.1726 | 42.8712 | -1.7356 | 0.18105 | -94.291 | 27.7404 | -94.291 | 27.7404 | 0.0029572 | 0.0089613
2020-01-08 | 1 | 0.02 | 933 | 930 | R Square | 0.99995 | Residual | 3 | 2509.3068 | 836.4356 |  |  | X Variable 1 | 47840.4643 | 186.2205 |  | 256.9022 | 1.3006e-07 | 47247.8274 | 48433.1012 | 47247.8274 | 48433.1012
2020-01-08 | 1 | 0.05 | 2377 | 2374 | Adjusted R Square | 0.99994 | Total | 4 | 55206195.2
2020-01-08 | 1 | 0.1 | 4710 | 4707 | Standard Error | 28.9212
2020-01-08 | 1 | 0.2 | 9555 | 9552 | Observations | 5
2020-01-08 | 2 | 0 | 5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.01 | 414 | 409 | Multiple R | 0.99997 | Regression | 1 | 45509099.5358 | 45509099.5358 | 44418.4174 | 2.3555e-07 | Intercept | -14.6169 | 21.2193 | 47.4479 | -0.68885 | 0.54041 | -82.1463 | 52.9125 | -82.1463 | 52.9125 | 0.0036047 | 0.010923
2020-01-08 | 2 | 0.02 | 866 | 861 | R Square | 0.99993 | Residual | 3 | 3073.6642 | 1024.5547 |  |  | X Variable 1 | 43437.0647 | 206.1004 |  | 210.7568 | 2.3555e-07 | 42781.1611 | 44092.9683 | 42781.1611 | 44092.9683
2020-01-08 | 2 | 0.05 | 2195 | 2190 | Adjusted R Square | 0.99991 | Total | 4 | 45512173.2
2020-01-08 | 2 | 0.1 | 4293 | 4288 | Standard Error | 32.0087
2020-01-08 | 2 | 0.2 | 8690 | 8685 | Observations | 5
2020-01-08 | 3 | 0 | 8 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 414 | 406 | Multiple R | 0.99992 | Regression | 1 | 42394098.8433 | 42394098.8433 | 18984.8218 | 8.429e-07 | Intercept | 4.7662 | 31.3266 | 70.0485 | 0.15214 | 0.88873 | -94.9292 | 104.4615 | -94.9292 | 104.4615 | 0.0055138 | 0.016708
2020-01-08 | 3 | 0.02 | 842 | 834 | R Square | 0.99984 | Residual | 3 | 6699.1567 | 2233.0522 |  |  | X Variable 1 | 41924.1294 | 304.2712 |  | 137.7854 | 8.429e-07 | 40955.8027 | 42892.456 | 40955.8027 | 42892.456
2020-01-08 | 3 | 0.05 | 2095 | 2087 | Adjusted R Square | 0.99979 | Total | 4 | 42400798
2020-01-08 | 3 | 0.1 | 4277 | 4269 | Standard Error | 47.2552
2020-01-08 | 3 | 0.2 | 8367 | 8359 | Observations | 5
2019-11-20 | 1 | 0 | 7 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.01 | 629.0139 | 622.0139 | Multiple R | 0.99995 | Regression | 1 | 85030459.2175 | 85030459.2175 | 27965.9048 | 4.7149e-07 | Intercept | 78.4873 | 36.5542 | 81.7376 | 2.1472 | 0.12103 | -37.8444 | 194.8191 | -37.8444 | 194.8191 | 0.0045429 | 0.013766
2019-11-20 | 1 | 0.02 | 1258.3054 | 1251.3054 | R Square | 0.99989 | Residual | 3 | 9121.5135 | 3040.5045 |  |  | X Variable 1 | 59374.3148 | 355.0456 |  | 167.2301 | 4.7149e-07 | 58244.4011 | 60504.2284 | 58244.4011 | 60504.2284
2019-11-20 | 1 | 0.05 | 3131.8435 | 3124.8435 | Adjusted R Square | 0.99986 | Total | 4 | 85039580.731
2019-11-20 | 1 | 0.1 | 6028.2724 | 6021.2724 | Standard Error | 55.1408
2019-11-20 | 1 | 0.2 | 11942.241 | 11935.241 | Observations | 5
2020-02-27 | 1 | 0 | 4.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.01 | 716.768 | 712.268 | Multiple R | 1 | Regression | 1 | 130705580.3247 | 130705580.3247 | 448652.2829 | 7.3384e-09 | Intercept | -34.089 | 11.315 | 25.3012 | -3.0127 | 0.057089 | -70.0985 | 1.9206 | -70.0985 | 1.9206 | 0.0011342 | 0.003437
2020-02-27 | 1 | 0.02 | 1451.3238 | 1446.8238 | R Square | 0.99999 | Residual | 3 | 873.9881 | 291.3294 |  |  | X Variable 1 | 73613.6604 | 109.9015 |  | 669.8151 | 7.3384e-09 | 73263.9049 | 73963.4159 | 73263.9049 | 73963.4159
2020-02-27 | 1 | 0.05 | 3625.21 | 3620.71 | Adjusted R Square | 0.99999 | Total | 4 | 130706454.3127
2020-02-27 | 1 | 0.1 | 7335.6332 | 7331.1332 | Standard Error | 17.0684
2020-02-27 | 1 | 0.2 | 14696.3111 | 14691.8111 | Observations | 5
2020-02-27 | 2 | 0 | 7.75 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | Standard Error | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.01 | 697.517 | 689.767 | Multiple R | 0.99996 | Regression | 1 | 107424162.7809 | 107424162.7809 | 42358.8383 | 2.5294e-07 | Intercept | 46.1604 | 33.3844 | 74.6497 | 1.3827 | 0.26071 | -60.0836 | 152.4044 | -60.0836 | 152.4044 | 0.0036913 | 0.011186
2020-02-27 | 2 | 0.02 | 1367.0655 | 1359.3155 | R Square | 0.99993 | Residual | 3 | 7608.1522 | 2536.0507 |  |  | X Variable 1 | 66736.3326 | 324.2577 |  | 205.8126 | 2.5294e-07 | 65704.3999 | 67768.2654 | 65704.3999 | 67768.2654
2020-02-27 | 2 | 0.05 | 3398.1542 | 3390.4042 | Adjusted R Square | 0.99991 | Total | 4 | 107431770.933
2020-02-27 | 2 | 0.1 | 6800.3691 | 6792.6191 | Standard Error | 50.3592
2020-02-27 | 2 | 0.2 | 13366.2523 | 13358.5023 | Observations | 5

## barium_Ba
Date | No_cal | Std_conc (ppb) | Ba (cps) | Ba (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 79.4537 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.05 | 1895.3261 | 1815.8724 | Multiple R | 0.99996 | Regression | 1 | 859884417.1637 | 859884417.1637 | 41637.5635 | 2.5954e-07 | Intercept | -130.0174 | 95.2668 | 213.0231 | -1.3648 | 0.26567 | -433.1989 | 173.1641 | -433.1989 | 173.1641 | 0.018616 | 0.056411
2020-01-08 | 1 | 0.1 | 3786.4439 | 3706.9902 | R Square | 0.99993 | Residual | 3 | 61954.9521 | 20651.6507 |  |  | X Variable 1 | 37762.5561 | 185.0626 |  | 204.0528 | 2.5954e-07 | 37173.6042 | 38351.508 | 37173.6042 | 38351.508
2020-01-08 | 1 | 0.25 | 9385.7736 | 9306.3199 | Adjusted R Square | 0.9999 | Total | 4 | 859946372.1157
2020-01-08 | 1 | 0.5 | 18618.0657 | 18538.612 | Standard Error | 143.7068
2020-01-08 | 1 | 1 | 37810.4289 | 37730.9752 | Observations | 5
2020-01-08 | 2 | 0 | 78.4485 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.05 | 1762.5353 | 1684.0868 | Multiple R | 0.99999 | Regression | 1 | 726553653.6654 | 726553653.6654 | 206180.4045 | 2.3555e-08 | Intercept | -72.9154 | 39.3527 | 87.9953 | -1.8529 | 0.16096 | -198.1533 | 52.3224 | -198.1533 | 52.3224 | 0.0083656 | 0.02535
2020-01-08 | 2 | 0.1 | 3479.3613 | 3400.9128 | R Square | 0.99999 | Residual | 3 | 10571.6204 | 3523.8735 |  |  | X Variable 1 | 34711.6445 | 76.4454 |  | 454.0709 | 2.3555e-08 | 34468.3611 | 34954.928 | 34468.3611 | 34954.928
2020-01-08 | 2 | 0.25 | 8712.5977 | 8634.1492 | Adjusted R Square | 0.99998 | Total | 4 | 726564225.2858
2020-01-08 | 2 | 0.5 | 17272.132 | 17193.6835 | Standard Error | 59.3622
2020-01-08 | 2 | 1 | 34753.1634 | 34674.7149 | Observations | 5
2020-01-08 | 3 | 0 | 92.4467 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.05 | 1708.7904 | 1616.3437 | Multiple R | 0.99999 | Regression | 1 | 674867465.823 | 674867465.823 | 109262.5539 | 6.1059e-08 | Intercept | -54.9397 | 52.1 | 116.4992 | -1.0545 | 0.36907 | -220.7452 | 110.8658 | -220.7452 | 110.8658 | 0.011492 | 0.034823
2020-01-08 | 3 | 0.1 | 3307.5658 | 3215.1191 | R Square | 0.99997 | Residual | 3 | 18529.7005 | 6176.5668 |  |  | X Variable 1 | 33454.1955 | 101.208 |  | 330.5489 | 6.1059e-08 | 33132.1064 | 33776.2845 | 33132.1064 | 33776.2845
2020-01-08 | 3 | 0.25 | 8511.0406 | 8418.5939 | Adjusted R Square | 0.99996 | Total | 4 | 674885995.5235
2020-01-08 | 3 | 0.5 | 16737.9874 | 16645.5407 | Standard Error | 78.5911
2020-01-08 | 3 | 1 | 33485.1224 | 33392.6757 | Observations | 5
2019-11-20 | 1 | 0 | 119.6792 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.05 | 2770.17 | 2650.4908 | Multiple R | 0.99999 | Regression | 1 | 1519627268.2785 | 1519627268.2785 | 162961.5543 | 3.3522e-08 | Intercept | 168.7653 | 64.0163 | 143.1447 | 2.6363 | 0.077903 | -34.963 | 372.4936 | -34.963 | 372.4936 | 0.0094098 | 0.028514
2019-11-20 | 1 | 0.1 | 5217.8709 | 5098.1917 | R Square | 0.99998 | Residual | 3 | 27975.1984 | 9325.0661 |  |  | X Variable 1 | 50200.7127 | 124.3562 |  | 403.685 | 3.3522e-08 | 49804.9559 | 50596.4695 | 49804.9559 | 50596.4695
2019-11-20 | 1 | 0.25 | 12919.7339 | 12800.0547 | Adjusted R Square | 0.99998 | Total | 4 | 1519655243.4768
2019-11-20 | 1 | 0.5 | 25484.0842 | 25364.405 | Standard Error | 96.5664
2019-11-20 | 1 | 1 | 50431.7178 | 50312.0385 | Observations | 5
2020-02-27 | 1 | 0 | 188.6583 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.05 | 2794.6898 | 2606.0315 | Multiple R | 0.99998 | Regression | 1 | 1801213040.3471 | 1801213040.3471 | 94360.9157 | 7.6079e-08 | Intercept | -225.9048 | 91.5906 | 204.8027 | -2.4665 | 0.090347 | -517.3868 | 65.5773 | -517.3868 | 65.5773 | 0.012366 | 0.037472
2020-02-27 | 1 | 0.1 | 5484.2008 | 5295.5425 | R Square | 0.99997 | Residual | 3 | 57265.6495 | 19088.5498 |  |  | X Variable 1 | 54654.2432 | 177.9212 |  | 307.1822 | 7.6079e-08 | 54088.0184 | 55220.468 | 54088.0184 | 55220.468
2020-02-27 | 1 | 0.25 | 13547.0738 | 13358.4155 | Adjusted R Square | 0.99996 | Total | 4 | 1801270305.9966
2020-02-27 | 1 | 0.5 | 27119.3689 | 26930.7106 | Standard Error | 138.1613
2020-02-27 | 1 | 1 | 54711.4965 | 54522.8382 | Observations | 5
2020-02-27 | 2 | 0 | 243.424 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE |  | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.05 | 2597.6551 | 2354.2311 | Multiple R | 0.99998 | Regression | 1 | 1507688925.4729 | 1507688925.4729 | 91489.328 | 7.9689e-08 | Intercept | -126.2283 | 85.1011 | 190.2917 | -1.4833 | 0.23464 | -397.0579 | 144.6012 | -397.0579 | 144.6012 | 0.012558 | 0.038056
2020-02-27 | 2 | 0.1 | 5160.3576 | 4916.9336 | R Square | 0.99997 | Residual | 3 | 49438.1899 | 16479.3966 |  |  | X Variable 1 | 50003.133 | 165.3149 |  | 302.472 | 7.9689e-08 | 49477.0272 | 50529.2388 | 49477.0272 | 50529.2388
2020-02-27 | 2 | 0.25 | 12714.0837 | 12470.6596 | Adjusted R Square | 0.99996 | Total | 4 | 1507738363.6629
2020-02-27 | 2 | 0.5 | 24934.9023 | 24691.4783 | Standard Error | 128.3721
2020-02-27 | 2 | 1 | 50184.9324 | 49941.5084 | Observations | 5

## tungsten_W
Date | No_cal | Std_conc (ppb) | W (cps) | W (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 6 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 204 | 198 | Multiple R | 0.99989 | Regression | 1 | 10512128.1758 | 10512128.1758 | 13784.4439 | 1.3623e-06 | Intercept | -23.6103 | 18.3069 | 40.9355 | -1.2897 | 0.28759 | -81.871 | 34.6505 | -81.871 | 34.6505 | 0.0064708 | 0.019608
2020-01-08 | 1 | 0.02 | 391 | 385 | R Square | 0.99978 | Residual | 3 | 2287.8242 | 762.6081 |  |  | X Variable 1 | 20876.4511 | 177.8124 |  | 117.4072 | 1.3623e-06 | 20310.5727 | 21442.3295 | 20310.5727 | 21442.3295
2020-01-08 | 1 | 0.05 | 1046 | 1040 | Adjusted R Square | 0.99971 | Total | 4 | 10514416
2020-01-08 | 1 | 0.1 | 2032 | 2026 | Standard Error | 27.6154
2020-01-08 | 1 | 0.2 | 4172 | 4166 | Observations | 5
2020-01-08 | 2 | 0 | 5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.01 | 209 | 204 | Multiple R | 0.99997 | Regression | 1 | 8549129.0881 | 8549129.0881 | 51696.7748 | 1.8761e-07 | Intercept | 11.5771 | 8.525 | 19.0624 | 1.358 | 0.26757 | -15.5532 | 38.7074 | -15.5532 | 38.7074 | 0.0033413 | 0.010125
2020-01-08 | 2 | 0.02 | 381 | 376 | R Square | 0.99994 | Residual | 3 | 496.1119 | 165.3706 |  |  | X Variable 1 | 18826.6169 | 82.8019 |  | 227.3692 | 1.8761e-07 | 18563.1042 | 19090.1297 | 18563.1042 | 19090.1297
2020-01-08 | 2 | 0.05 | 956 | 951 | Adjusted R Square | 0.99992 | Total | 4 | 8549625.2
2020-01-08 | 2 | 0.1 | 1916 | 1911 | Standard Error | 12.8597
2020-01-08 | 2 | 0.2 | 3775 | 3770 | Observations | 5
2020-01-08 | 3 | 0 | 5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 193 | 188 | Multiple R | 0.99976 | Regression | 1 | 8182267.2265 | 8182267.2265 | 6339.61 | 4.3665e-06 | Intercept | 13.6136 | 23.8161 | 53.2543 | 0.57161 | 0.6076 | -62.1797 | 89.4069 | -62.1797 | 89.4069 | 0.0095416 | 0.028914
2020-01-08 | 3 | 0.02 | 363 | 358 | R Square | 0.99953 | Residual | 3 | 3871.9735 | 1290.6578 |  |  | X Variable 1 | 18418.2421 | 231.322 |  | 79.6217 | 4.3665e-06 | 17682.0723 | 19154.4119 | 17682.0723 | 19154.4119
2020-01-08 | 3 | 0.05 | 993 | 988 | Adjusted R Square | 0.99937 | Total | 4 | 8186139.2
2020-01-08 | 3 | 0.1 | 1842 | 1837 | Standard Error | 35.9257
2020-01-08 | 3 | 0.2 | 3701 | 3696 | Observations | 5
2019-11-20 | 1 | 0 | 1.75 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.01 | 278.7524 | 277.0024 | Multiple R | 0.99994 | Regression | 1 | 17202512.3625 | 17202512.3625 | 24520.4017 | 5.7427e-07 | Intercept | 25.0193 | 17.5589 | 39.2628 | 1.4249 | 0.24941 | -30.8608 | 80.8995 | -30.8608 | 80.8995 | 0.0048516 | 0.014702
2019-11-20 | 1 | 0.02 | 549.0102 | 547.2602 | R Square | 0.99988 | Residual | 3 | 2104.6775 | 701.5592 |  |  | X Variable 1 | 26705.9043 | 170.5468 |  | 156.5899 | 5.7427e-07 | 26163.1483 | 27248.6602 | 26163.1483 | 27248.6602
2019-11-20 | 1 | 0.05 | 1402.3185 | 1400.5685 | Adjusted R Square | 0.99984 | Total | 4 | 17204617.04
2019-11-20 | 1 | 0.1 | 2687.0021 | 2685.2521 | Standard Error | 26.487
2019-11-20 | 1 | 0.2 | 5365.0071 | 5363.2571 | Observations | 5
2020-02-27 | 1 | 0 | 1.4997 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.01 | 248.0018 | 246.5022 | Multiple R | 0.99994 | Regression | 1 | 14099922.6423 | 14099922.6423 | 26137.0059 | 5.2183e-07 | Intercept | 7.443 | 15.3973 | 34.4294 | 0.4834 | 0.6619 | -41.5581 | 56.4441 | -41.5581 | 56.4441 | 0.0046992 | 0.01424
2020-02-27 | 1 | 0.02 | 490.0084 | 488.5087 | R Square | 0.99989 | Residual | 3 | 1618.3861 | 539.462 |  |  | X Variable 1 | 24177.9633 | 149.5519 |  | 161.6694 | 5.2183e-07 | 23702.0225 | 24653.904 | 23702.0225 | 24653.904
2020-02-27 | 1 | 0.05 | 1243.8042 | 1242.3045 | Adjusted R Square | 0.99985 | Total | 4 | 14101541.0284
2020-02-27 | 1 | 0.1 | 2397.4513 | 2395.9516 | Standard Error | 23.2263
2020-02-27 | 1 | 0.2 | 4853.0737 | 4851.574 | Observations | 5
2020-02-27 | 2 | 0 | 1.75 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.01 | 231.5016 | 229.7516 | Multiple R | 0.99981 | Regression | 1 | 10831307.6005 | 10831307.6005 | 7742.597 | 3.2355e-06 | Intercept | 33.3534 | 24.7948 | 55.443 | 1.3452 | 0.27122 | -45.5549 | 112.2616 | -45.5549 | 112.2616 | 0.0086339 | 0.026163
2020-02-27 | 2 | 0.02 | 443.2569 | 441.5069 | R Square | 0.99961 | Residual | 3 | 4196.7731 | 1398.9244 |  |  | X Variable 1 | 21191.0167 | 240.8288 |  | 87.992 | 3.2355e-06 | 20424.5919 | 21957.4415 | 20424.5919 | 21957.4415
2020-02-27 | 2 | 0.05 | 1096.2921 | 1094.5421 | Adjusted R Square | 0.99948 | Total | 4 | 10835504.3736
2020-02-27 | 2 | 0.1 | 2209.421 | 2207.671 | Standard Error | 37.4022
2020-02-27 | 2 | 0.2 | 4247.6315 | 4245.8815 | Observations | 5

## thallium_Tl
Date | No_cal | Std_conc (ppb) | Tl (cps) | Tl (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 2 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.05 | 2900.5446 | 2898.5446 | Multiple R | 1 | Regression | 1 | 2175318926.0904 | 2175318926.0904 | 335079.8191 | 1.137e-08 | Intercept | -199.4455 | 53.4136 | 119.4364 | -3.734 | 0.033483 | -369.4313 | -29.4597 | -369.4313 | -29.4597 | 0.0065622 | 0.019885
2020-01-08 | 1 | 0.1 | 5801.6781 | 5799.6781 | R Square | 0.99999 | Residual | 3 | 19475.8276 | 6491.9425 |  |  | X Variable 1 | 60062.4181 | 103.7597 |  | 578.8608 | 1.137e-08 | 59732.2085 | 60392.6278 | 59732.2085 | 60392.6278
2020-01-08 | 1 | 0.25 | 14726.8383 | 14724.8383 | Adjusted R Square | 0.99999 | Total | 4 | 2175338401.9179
2020-01-08 | 1 | 0.5 | 29803.306 | 29801.306 | Standard Error | 80.5726
2020-01-08 | 1 | 1 | 59899 | 59897 | Observations | 5
2020-01-08 | 2 | 0 | 12.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.05 | 2622.4907 | 2609.9907 | Multiple R | 0.99998 | Regression | 1 | 1734192091.3278 | 1734192091.3278 | 76466.2479 | 1.0429e-07 | Intercept | -118.578 | 99.8339 | 223.2354 | -1.1878 | 0.3204 | -436.294 | 199.138 | -436.294 | 199.138 | 0.013737 | 0.041627
2020-01-08 | 2 | 0.1 | 5246.7135 | 5234.2135 | R Square | 0.99996 | Residual | 3 | 68037.5514 | 22679.1838 |  |  | X Variable 1 | 53627.7955 | 193.9345 |  | 276.5253 | 1.0429e-07 | 53010.6094 | 54244.9816 | 53010.6094 | 54244.9816
2020-01-08 | 2 | 0.25 | 13400.2828 | 13387.7828 | Adjusted R Square | 0.99995 | Total | 4 | 1734260128.8792
2020-01-08 | 2 | 0.5 | 26487.0404 | 26474.5404 | Standard Error | 150.5961
2020-01-08 | 2 | 1 | 53605.894 | 53593.394 | Observations | 5
2020-01-08 | 3 | 0 | 23 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.05 | 2457.9616 | 2434.9616 | Multiple R | 0.99999 | Regression | 1 | 1582210505.8294 | 1582210505.8294 | 103263.0628 | 6.6457e-08 | Intercept | -146.5361 | 82.0586 | 183.4885 | -1.7857 | 0.17212 | -407.683 | 114.6109 | -407.683 | 114.6109 | 0.011821 | 0.035821
2020-01-08 | 3 | 0.1 | 5068.6491 | 5045.6491 | R Square | 0.99997 | Residual | 3 | 45966.4026 | 15322.1342 |  |  | X Variable 1 | 51223.9985 | 159.4046 |  | 321.3457 | 6.6457e-08 | 50716.7017 | 51731.2952 | 50716.7017 | 51731.2952
2020-01-08 | 3 | 0.25 | 12695.8908 | 12672.8908 | Adjusted R Square | 0.99996 | Total | 4 | 1582256472.232
2020-01-08 | 3 | 0.5 | 25304.1431 | 25281.1431 | Standard Error | 123.7826
2020-01-08 | 3 | 1 | 51181.2723 | 51158.2722 | Observations | 5
2019-11-20 | 1 | 0 | 2 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.05 | 3374.1485 | 3372.1485 | Multiple R | 0.99989 | Regression | 1 | 2322092580.4571 | 2322092580.4571 | 13264.2522 | 1.4432e-06 | Intercept | 554.7115 | 277.372 | 620.2227 | 1.9999 | 0.13934 | -328.01 | 1437.4331 | -328.01 | 1437.4331 | 0.032982 | 0.099946
2019-11-20 | 1 | 0.1 | 6627.2871 | 6625.2871 | R Square | 0.99977 | Residual | 3 | 525191.8937 | 175063.9646 |  |  | X Variable 1 | 62055.6186 | 538.8151 |  | 115.1705 | 1.4432e-06 | 60340.8686 | 63770.3686 | 60340.8686 | 63770.3686
2019-11-20 | 1 | 0.25 | 16217.451 | 16215.451 | Adjusted R Square | 0.9997 | Total | 4 | 2322617772.3508
2019-11-20 | 1 | 0.5 | 32149.3977 | 32147.3977 | Standard Error | 418.4065
2019-11-20 | 1 | 1 | 62320.9489 | 62318.9489 | Observations | 5
2020-02-27 | 1 | 0 | 1.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.05 | 3059.5777 | 3058.3277 | Multiple R | 0.88872 | Regression | 1 | 2790536758.6977 | 2790536758.6977 | 11.2738 | 0.04381 | Intercept | 3829.81 | 10429.7258 | 23321.5758 | 0.3672 | 0.73784 | -29362.2323 | 37021.8523 | -29362.2323 | 37021.8523 | 1.1313 | 3.4283
2020-02-27 | 1 | 0.1 | 6173.5839 | 6172.3339 | R Square | 0.78982 | Residual | 3 | 742571836.2236 | 247523945.4079 |  |  | X Variable 1 | 68027.6107 | 20260.4904 |  | 3.3576 | 0.04381 | 3549.6879 | 132505.5336 | 3549.6879 | 132505.5336
2020-02-27 | 1 | 0.25 | 15471.3763 | 15470.1263 | Adjusted R Square | 0.71977 | Total | 4 | 3533108594.9213
2020-02-27 | 1 | 0.5 | 61851.6112 | 61850.3612 | Standard Error | 15732.8937
2020-02-27 | 1 | 1 | 61851.6112 | 61850.3612 | Observations | 5
2020-02-27 | 2 | 0 | 1 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | Standard Error | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.05 | 2598.4864 | 2597.4864 | Multiple R | 0.99998 | Regression | 1 | 1630103491.1785 | 1630103491.1785 | 95733.9341 | 7.4448e-08 | Intercept | 147.8178 | 86.5045 | 193.43 | 1.7088 | 0.18602 | -127.4783 | 423.1139 | -127.4783 | 423.1139 | 0.012277 | 0.037203
2020-02-27 | 2 | 0.1 | 5353.0029 | 5352.0029 | R Square | 0.99997 | Residual | 3 | 51082.3097 | 17027.4366 |  |  | X Variable 1 | 51993.4855 | 168.0413 |  | 309.409 | 7.4448e-08 | 51458.7032 | 52528.2679 | 51458.7032 | 52528.2679
2020-02-27 | 2 | 0.25 | 13299.9388 | 13298.9388 | Adjusted R Square | 0.99996 | Total | 4 | 1630154573.4882
2020-02-27 | 2 | 0.5 | 26192.741 | 26191.741 | Standard Error | 130.4892
2020-02-27 | 2 | 1 | 52087.5423 | 52086.5423 | Observations | 5

## lead_Pb
Date | No_cal | Std_conc (ppb) | Pb (cps) | Pb (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 193.2505 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.05 | 3902.7108 | 3709.4603 | Multiple R | 0.99994 | Regression | 1 | 3764531560.1857 | 3764531560.1857 | 25683.498 | 5.3571e-07 | Intercept | -440.9452 | 253.8006 | 567.5154 | -1.7374 | 0.18072 | -1248.652 | 366.7617 | -1248.652 | 366.7617 | 0.023703 | 0.071826
2020-01-08 | 1 | 0.1 | 7860.1001 | 7666.8496 | R Square | 0.99988 | Residual | 3 | 439721.8277 | 146573.9426 |  |  | X Variable 1 | 79012.6841 | 493.0259 |  | 160.2607 | 5.3571e-07 | 77443.6556 | 80581.7125 | 77443.6556 | 80581.7125
2020-01-08 | 1 | 0.25 | 19345.3995 | 19152.149 | Adjusted R Square | 0.99984 | Total | 4 | 3764971282.0134
2020-01-08 | 1 | 0.5 | 38747.4499 | 38554.1994 | Standard Error | 382.8498
2020-01-08 | 1 | 1 | 79029.9662 | 78836.7157 | Observations | 5
2020-01-08 | 2 | 0 | 183.0004 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.05 | 3543.6725 | 3360.6721 | Multiple R | 0.99999 | Regression | 1 | 2926984910.2213 | 2926984910.2213 | 165274.0687 | 3.2821e-08 | Intercept | -45.2008 | 88.221 | 197.2682 | -0.51236 | 0.64374 | -325.9595 | 235.5579 | -325.9595 | 235.5579 | 0.0093437 | 0.028314
2020-01-08 | 2 | 0.1 | 7028.1755 | 6845.175 | R Square | 0.99998 | Residual | 3 | 53129.6579 | 17709.886 |  |  | X Variable 1 | 69670.9265 | 171.3757 |  | 406.5391 | 3.2821e-08 | 69125.5325 | 70216.3204 | 69125.5325 | 70216.3204
2020-01-08 | 2 | 0.25 | 17755.5974 | 17572.597 | Adjusted R Square | 0.99998 | Total | 4 | 2927038039.8792
2020-01-08 | 2 | 0.5 | 34958.8331 | 34775.8327 | Standard Error | 133.0785
2020-01-08 | 2 | 1 | 69777.4797 | 69594.4793 | Observations | 5
2020-01-08 | 3 | 0 | 223.7507 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.05 | 3409.9088 | 3186.1581 | Multiple R | 0.99995 | Regression | 1 | 2721808347.4859 | 2721808347.4859 | 27494.304 | 4.8367e-07 | Intercept | -218.8476 | 208.5795 | 466.398 | -1.0492 | 0.37114 | -882.6408 | 444.9455 | -882.6408 | 444.9455 | 0.022909 | 0.06942
2020-01-08 | 3 | 0.1 | 6762.6276 | 6538.8769 | R Square | 0.99989 | Residual | 3 | 296986.0607 | 98995.3536 |  |  | X Variable 1 | 67184.6586 | 405.1807 |  | 165.8141 | 4.8367e-07 | 65895.1928 | 68474.1244 | 65895.1928 | 68474.1244
2020-01-08 | 3 | 0.25 | 17008.2376 | 16784.487 | Adjusted R Square | 0.99985 | Total | 4 | 2722105333.5466
2020-01-08 | 3 | 0.5 | 33128.8433 | 32905.0927 | Standard Error | 314.6353
2020-01-08 | 3 | 1 | 67365.7491 | 67141.9985 | Observations | 5
2019-11-20 | 1 | 0 | 88.2501 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.05 | 4430.7677 | 4342.5176 | Multiple R | 0.99994 | Regression | 1 | 3851358351.5019 | 3851358351.5019 | 24142.6078 | 5.878e-07 | Intercept | 687.2459 | 264.7763 | 592.0579 | 2.5956 | 0.080684 | -155.3905 | 1529.8824 | -155.3905 | 1529.8824 | 0.024447 | 0.074083
2019-11-20 | 1 | 0.1 | 8596.0116 | 8507.7615 | R Square | 0.99988 | Residual | 3 | 478576.0981 | 159525.366 |  |  | X Variable 1 | 79918.6812 | 514.347 |  | 155.3789 | 5.878e-07 | 78281.7995 | 81555.563 | 78281.7995 | 81555.563
2019-11-20 | 1 | 0.25 | 21269.1919 | 21180.9418 | Adjusted R Square | 0.99983 | Total | 4 | 3851836927.6
2019-11-20 | 1 | 0.5 | 40919.4913 | 40831.2412 | Standard Error | 399.4063
2019-11-20 | 1 | 1 | 80507.5119 | 80419.2618 | Observations | 5
2020-02-27 | 1 | 0 | 148.2503 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.05 | 4118.7321 | 3970.4818 | Multiple R | 0.99997 | Regression | 1 | 4088175336.8644 | 4088175336.8644 | 58055.2487 | 1.5765e-07 | Intercept | -245.2525 | 175.9172 | 393.3628 | -1.3941 | 0.25759 | -805.0995 | 314.5945 | -805.0995 | 314.5945 | 0.015765 | 0.047774
2020-02-27 | 1 | 0.1 | 8219.4296 | 8071.1794 | R Square | 0.99995 | Residual | 3 | 211256.1101 | 70418.7034 |  |  | X Variable 1 | 82339.0972 | 341.7318 |  | 240.9466 | 1.5765e-07 | 81251.5542 | 83426.6402 | 81251.5542 | 83426.6402
2020-02-27 | 1 | 0.25 | 20531.5483 | 20383.2981 | Adjusted R Square | 0.99993 | Total | 4 | 4088386592.9745
2020-02-27 | 1 | 0.5 | 40670.6613 | 40522.411 | Standard Error | 265.3652
2020-02-27 | 1 | 1 | 82418.9023 | 82270.652 | Observations | 5
2020-02-27 | 2 | 0 | 246.0015 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE |  | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.05 | 3575.4246 | 3329.4231 | Multiple R | 0.99998 | Regression | 1 | 2803399106.5517 | 2803399106.5517 | 80274.0565 | 9.6959e-08 | Intercept | 48.5246 | 123.8851 | 277.0155 | 0.39169 | 0.72145 | -345.7331 | 442.7823 | -345.7331 | 442.7823 | 0.013407 | 0.040628
2020-02-27 | 2 | 0.1 | 7030.1823 | 6784.1809 | R Square | 0.99996 | Residual | 3 | 104768.5602 | 34922.8534 |  |  | X Variable 1 | 68184.2094 | 240.6557 |  | 283.3268 | 9.6959e-08 | 67418.3355 | 68950.0833 | 67418.3355 | 68950.0833
2020-02-27 | 2 | 0.25 | 17621.5312 | 17375.5297 | Adjusted R Square | 0.99995 | Total | 4 | 2803503875.1119
2020-02-27 | 2 | 0.5 | 34357.9708 | 34111.9693 | Standard Error | 186.8766
2020-02-27 | 2 | 1 | 68437.5194 | 68191.5179 | Observations | 5

## uranium_U
Date | No_cal | Std_conc (ppb) | U (cps) | U (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 1 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 754.5199 | 753.5199 | Multiple R | 0.99998 | Regression | 1 | 154727899.0511 | 154727899.0511 | 92252.1227 | 7.8703e-08 | Intercept | -75.0856 | 27.1494 | 60.7079 | -2.7656 | 0.069822 | -161.4871 | 11.316 | -161.4871 | 11.316 | 0.0025013 | 0.0075797
2020-01-08 | 1 | 0.02 | 1547.5839 | 1546.5839 | R Square | 0.99997 | Residual | 3 | 5031.6858 | 1677.2286 |  |  | X Variable 1 | 80093.2031 | 263.6984 |  | 303.7303 | 7.8703e-08 | 79253.9972 | 80932.4091 | 79253.9972 | 80932.4091
2020-01-08 | 1 | 0.05 | 3869.2741 | 3868.2741 | Adjusted R Square | 0.99996 | Total | 4 | 154732930.7369
2020-01-08 | 1 | 0.1 | 7938.956 | 7937.956 | Standard Error | 40.954
2020-01-08 | 1 | 0.2 | 15954.6553 | 15953.6553 | Observations | 5
2020-01-08 | 2 | 0 | 2 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.01 | 696.767 | 694.767 | Multiple R | 0.99997 | Regression | 1 | 121301278.3284 | 121301278.3284 | 54101.3504 | 1.7524e-07 | Intercept | -12.5523 | 31.3901 | 70.1904 | -0.39988 | 0.71601 | -112.4497 | 87.3451 | -112.4497 | 87.3451 | 0.0032662 | 0.0098977
2020-01-08 | 2 | 0.02 | 1421.3208 | 1419.3208 | R Square | 0.99994 | Residual | 3 | 6726.3355 | 2242.1118 |  |  | X Variable 1 | 70915.9683 | 304.8878 |  | 232.597 | 1.7524e-07 | 69945.6794 | 71886.2572 | 69945.6794 | 71886.2572
2020-01-08 | 2 | 0.05 | 3567.6955 | 3565.6955 | Adjusted R Square | 0.99993 | Total | 4 | 121308004.664
2020-01-08 | 2 | 0.1 | 7011.4705 | 7009.4705 | Standard Error | 47.3509
2020-01-08 | 2 | 0.2 | 14198.0527 | 14196.0527 | Observations | 5
2020-01-08 | 3 | 0 | 6.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 657.0151 | 650.7651 | Multiple R | 0.99998 | Regression | 1 | 107496779.7771 | 107496779.7771 | 83492.9048 | 9.1407e-08 | Intercept | 6.6823 | 23.7869 | 53.189 | 0.28092 | 0.79703 | -69.0181 | 82.3827 | -69.0181 | 82.3827 | 0.0026292 | 0.0079673
2020-01-08 | 3 | 0.02 | 1334.5624 | 1328.3124 | R Square | 0.99996 | Residual | 3 | 3862.488 | 1287.496 |  |  | X Variable 1 | 66758.8852 | 231.0385 |  | 288.9514 | 9.1407e-08 | 66023.6176 | 67494.1527 | 66023.6176 | 67494.1527
2020-01-08 | 3 | 0.05 | 3364.8963 | 3358.6463 | Adjusted R Square | 0.99995 | Total | 4 | 107500642.2651
2020-01-08 | 3 | 0.1 | 6736.8383 | 6730.5883 | Standard Error | 35.8817
2020-01-08 | 3 | 0.2 | 13339.7257 | 13333.4757 | Observations | 5
2019-11-20 | 1 | 0 | 0.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.01 | 824.7738 | 824.2738 | Multiple R | 0.99993 | Regression | 1 | 129633187.0447 | 129633187.0447 | 22791.8225 | 6.4082e-07 | Intercept | 146.7618 | 49.9957 | 111.7938 | 2.9355 | 0.060731 | -12.3469 | 305.8705 | -12.3469 | 305.8705 | 0.0050323 | 0.015249
2019-11-20 | 1 | 0.02 | 1583.5878 | 1583.0878 | R Square | 0.99987 | Residual | 3 | 17063.1182 | 5687.7061 |  |  | X Variable 1 | 73311.0513 | 485.6014 |  | 150.9696 | 6.4082e-07 | 71765.651 | 74856.4516 | 71765.651 | 74856.4516
2019-11-20 | 1 | 0.05 | 3858.021 | 3857.521 | Adjusted R Square | 0.99982 | Total | 4 | 129650250.1629
2019-11-20 | 1 | 0.1 | 7570.0054 | 7569.5054 | Standard Error | 75.4169
2019-11-20 | 1 | 0.2 | 14758.1207 | 14757.6207 | Observations | 5
2020-02-27 | 1 | 0 | 0.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.01 | 738.5191 | 738.2691 | Multiple R | 0.99996 | Regression | 1 | 139460921.6339 | 139460921.6339 | 33859.5177 | 3.5392e-07 | Intercept | -63.5109 | 42.5452 | 95.1339 | -1.4928 | 0.23232 | -198.9086 | 71.8868 | -198.9086 | 71.8868 | 0.0041287 | 0.012511
2020-02-27 | 1 | 0.02 | 1451.0738 | 1450.8238 | R Square | 0.99991 | Residual | 3 | 12356.4301 | 4118.81 |  |  | X Variable 1 | 76039.2131 | 413.2351 |  | 184.0096 | 3.5392e-07 | 74724.1145 | 77354.3117 | 74724.1145 | 77354.3117
2020-02-27 | 1 | 0.05 | 3755.994 | 3755.744 | Adjusted R Square | 0.99988 | Total | 4 | 139473278.064
2020-02-27 | 1 | 0.1 | 7447.691 | 7447.441 | Standard Error | 64.178
2020-02-27 | 1 | 0.2 | 15185.3183 | 15185.0683 | Observations | 5
2020-02-27 | 2 | 0 | 0.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.01 | 637.7642 | 637.5142 | Multiple R | 0.99997 | Regression | 1 | 95431673.6576 | 95431673.6576 | 52643.0923 | 1.8257e-07 | Intercept | 13.9457 | 28.2254 | 63.1139 | 0.49408 | 0.65516 | -75.8801 | 103.7714 | -75.8801 | 103.7714 | 0.0033112 | 0.010034
2020-02-27 | 2 | 0.02 | 1247.0544 | 1246.8044 | R Square | 0.99994 | Residual | 3 | 5438.4157 | 1812.8052 |  |  | X Variable 1 | 62901.0099 | 274.1492 |  | 229.4408 | 1.8257e-07 | 62028.5449 | 63773.4749 | 62028.5449 | 63773.4749
2020-02-27 | 2 | 0.05 | 3219.1128 | 3218.8628 | Adjusted R Square | 0.99992 | Total | 4 | 95437112.0733
2020-02-27 | 2 | 0.1 | 6270.1263 | 6269.8763 | Standard Error | 42.5771
2020-02-27 | 2 | 0.2 | 12599.3044 | 12599.0544 | Observations | 5

## manganese_Mn
Date | No_cal | Std_conc (ppb) | Mn (cps) | Mn (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 150.5008 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 1484.0771 | 1333.5763 | Multiple R | 0.99975 | Regression | 1 | 320994484.9942 | 320994484.9942 | 6104.9988 | 4.6205e-06 | Intercept | -74.6269 | 152.0093 | 339.9032 | -0.49094 | 0.65714 | -558.3885 | 409.1346 | -558.3885 | 409.1346 | 0.0097232 | 0.029464
2020-01-08 | 1 | 0.02 | 2383.6989 | 2233.1981 | R Square | 0.99951 | Residual | 3 | 157736.8781 | 52578.9594 |  |  | X Variable 1 | 115361.2979 | 1476.4452 |  | 78.1345 | 4.6205e-06 | 110662.5903 | 120060.0056 | 110662.5903 | 120060.0056
2020-01-08 | 1 | 0.05 | 5650.3674 | 5499.8666 | Adjusted R Square | 0.99935 | Total | 4 | 321152221.8722
2020-01-08 | 1 | 0.1 | 11417.5608 | 11267.06 | Standard Error | 229.301
2020-01-08 | 1 | 0.2 | 23280.9583 | 23130.4575 | Observations | 5
2020-01-08 | 2 | 0 | 153.7508 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.01 | 1292.8086 | 1139.0577 | Multiple R | 0.99991 | Regression | 1 | 258216944.4181 | 258216944.4181 | 15978.2102 | 1.0916e-06 | Intercept | -63.256 | 84.2738 | 188.4419 | -0.7506 | 0.5074 | -331.4528 | 204.9409 | -331.4528 | 204.9409 | 0.0060102 | 0.018213
2020-01-08 | 2 | 0.02 | 2093.6535 | 1939.9027 | R Square | 0.99981 | Residual | 3 | 48481.7024 | 16160.5675 |  |  | X Variable 1 | 103467.4431 | 818.5395 |  | 126.4049 | 1.0916e-06 | 100862.4851 | 106072.4012 | 100862.4851 | 106072.4012
2020-01-08 | 2 | 0.05 | 5183.4404 | 5029.6895 | Adjusted R Square | 0.99975 | Total | 4 | 258265426.1205
2020-01-08 | 2 | 0.1 | 10358.507 | 10204.7562 | Standard Error | 127.1242
2020-01-08 | 2 | 0.2 | 20841.6933 | 20687.9425 | Observations | 5
2020-01-08 | 3 | 0 | 144.0007 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 837.5246 | 693.5238 | Multiple R | 0.99978 | Regression | 1 | 93864986.3261 | 93864986.3261 | 6781.3959 | 3.9469e-06 | Intercept | -38.5426 | 77.9932 | 174.398 | -0.49418 | 0.6551 | -286.7516 | 209.6665 | -286.7516 | 209.6665 | 0.0092256 | 0.027956
2020-01-08 | 3 | 0.02 | 1323.0614 | 1179.0607 | R Square | 0.99956 | Residual | 3 | 41524.6304 | 13841.5435 |  |  | X Variable 1 | 62382.5551 | 757.5366 |  | 82.3492 | 3.9469e-06 | 59971.7356 | 64793.3746 | 59971.7356 | 64793.3746
2020-01-08 | 3 | 0.05 | 3225.6154 | 3081.6146 | Adjusted R Square | 0.99941 | Total | 4 | 93906510.9565
2020-01-08 | 3 | 0.1 | 6190.598 | 6046.5973 | Standard Error | 117.6501
2020-01-08 | 3 | 0.2 | 12655.8624 | 12511.8617 | Observations | 5
2019-11-20 | 1 | 0 | 542.5103 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.01 | 772.0209 | 229.5106 | Multiple R | 0.99918 | Regression | 1 | 129681390.5278 | 129681390.5278 | 1834.4062 | 2.8014e-05 | Intercept | -644.2795 | 176.2606 | 394.1308 | -3.6553 | 0.035362 | -1205.2195 | -83.3394 | -1205.2195 | -83.3394 | 0.017738 | 0.053751
2019-11-20 | 1 | 0.02 | 1397.3184 | 854.8081 | R Square | 0.99837 | Residual | 3 | 212081.8031 | 70693.9344 |  |  | X Variable 1 | 73324.6802 | 1711.9947 |  | 42.83 | 2.8014e-05 | 67876.3489 | 78773.0115 | 67876.3489 | 78773.0115
2019-11-20 | 1 | 0.05 | 3612.2067 | 3069.6964 | Adjusted R Square | 0.99782 | Total | 4 | 129893472.3309
2019-11-20 | 1 | 0.1 | 6833.3848 | 6290.8745 | Standard Error | 265.8833
2019-11-20 | 1 | 0.2 | 14739.6019 | 14197.0916 | Observations | 5
2020-02-27 | 1 | 0 | 209.5015 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.01 | 831.5242 | 622.0227 | Multiple R | 0.99984 | Regression | 1 | 148171747.2577 | 148171747.2577 | 9443.1327 | 2.4023e-06 | Intercept | -226.3061 | 83.0403 | 185.6838 | -2.7253 | 0.072229 | -490.5774 | 37.9652 | -490.5774 | 37.9652 | 0.007818 | 0.023691
2020-02-27 | 1 | 0.02 | 1579.5874 | 1370.0859 | R Square | 0.99968 | Residual | 3 | 47072.8575 | 15690.9525 |  |  | X Variable 1 | 78377.9769 | 806.5587 |  | 97.1758 | 2.4023e-06 | 75811.147 | 80944.8068 | 75811.147 | 80944.8068
2020-02-27 | 1 | 0.05 | 3910.0353 | 3700.5337 | Adjusted R Square | 0.99958 | Total | 4 | 148218820.1152
2020-02-27 | 1 | 0.1 | 7634.5402 | 7425.0386 | Standard Error | 125.2635
2020-02-27 | 1 | 0.2 | 15743.9214 | 15534.4198 | Observations | 5
2020-02-27 | 2 | 0 | 403.0057 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | Standard Error | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.01 | 808.7729 | 405.7673 | Multiple R | 0.99996 | Regression | 1 | 150931390.8616 | 150931390.8616 | 36948.1224 | 3.1048e-07 | Intercept | -460.1998 | 42.3699 | 94.7421 | -10.8615 | 0.00167 | -595.0399 | -325.3597 | -595.0399 | -325.3597 | 0.0039524 | 0.011977
2020-02-27 | 2 | 0.02 | 1480.3268 | 1077.3211 | R Square | 0.99992 | Residual | 3 | 12254.8629 | 4084.9543 |  |  | X Variable 1 | 79104.4901 | 411.5333 |  | 192.2189 | 3.1048e-07 | 77794.8076 | 80414.1726 | 77794.8076 | 80414.1726
2020-02-27 | 2 | 0.05 | 3897.0318 | 3494.0261 | Adjusted R Square | 0.99989 | Total | 4 | 150943645.7245
2020-02-27 | 2 | 0.1 | 7792.6249 | 7389.6192 | Standard Error | 63.9136
2020-02-27 | 2 | 0.2 | 15794.9794 | 15391.9737 | Observations | 5

## selenium_Se
Date | No_cal | Std_conc (ppb) | Se (cps) | Se (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 183.2512 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.25 | 839.2747 | 656.0235 | Multiple R | 0.99998 | Regression | 1 | 110898037.6618 | 110898037.6618 | 70227.3594 | 1.1849e-07 | Intercept | -59.9678 | 26.3435 | 58.9058 | -2.2764 | 0.10729 | -143.8045 | 23.8689 | -143.8045 | 23.8689 | 0.07167 | 0.21718
2020-01-08 | 1 | 0.5 | 1488.8276 | 1305.5764 | R Square | 0.99996 | Residual | 3 | 4737.3861 | 1579.1287 |  |  | X Variable 1 | 2712.2722 | 10.2348 |  | 265.0045 | 1.1849e-07 | 2679.7004 | 2744.8439 | 2679.7004 | 2744.8439
2020-01-08 | 1 | 1.25 | 3479.6739 | 3296.4227 | Adjusted R Square | 0.99994 | Total | 4 | 110902775.0479
2020-01-08 | 1 | 2.5 | 6865.8996 | 6682.6484 | Standard Error | 39.7383
2020-01-08 | 1 | 5 | 13709.3269 | 13526.0757 | Observations | 5
2020-01-08 | 2 | 0 | 180.5011 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.25 | 758.2701 | 577.769 | Multiple R | 0.99989 | Regression | 1 | 80751915.4654 | 80751915.4654 | 13202.8062 | 1.4533e-06 | Intercept | -44.5207 | 51.8451 | 115.9291 | -0.85873 | 0.45363 | -209.5149 | 120.4735 | -209.5149 | 120.4735 | 0.16529 | 0.50089
2020-01-08 | 2 | 0.5 | 1284.8078 | 1104.3066 | R Square | 0.99977 | Residual | 3 | 18348.8073 | 6116.2691 |  |  | X Variable 1 | 2314.4497 | 20.1426 |  | 114.9035 | 1.4533e-06 | 2250.3471 | 2378.5524 | 2250.3471 | 2378.5524
2020-01-08 | 2 | 1.25 | 3060.0779 | 2879.5768 | Adjusted R Square | 0.9997 | Total | 4 | 80770264.2727
2020-01-08 | 2 | 2.5 | 5807.6806 | 5627.1794 | Standard Error | 78.2066
2020-01-08 | 2 | 5 | 11756.3384 | 11575.8372 | Observations | 5
2020-01-08 | 3 | 0 | 190.5013 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.25 | 691.0167 | 500.5155 | Multiple R | 0.99997 | Regression | 1 | 66071888.3534 | 66071888.3534 | 54861.649 | 1.7161e-07 | Intercept | -34.6524 | 23.0058 | 51.4426 | -1.5062 | 0.22909 | -107.8673 | 38.5625 | -107.8673 | 38.5625 | 0.081088 | 0.24572
2020-01-08 | 3 | 0.5 | 1157.5469 | 967.0457 | R Square | 0.99995 | Residual | 3 | 3613.0096 | 1204.3365 |  |  | X Variable 1 | 2093.5325 | 8.9381 |  | 234.2256 | 1.7161e-07 | 2065.0874 | 2121.9775 | 2065.0874 | 2121.9775
2020-01-08 | 3 | 1.25 | 2810.0264 | 2619.5252 | Adjusted R Square | 0.99993 | Total | 4 | 66075501.3631
2020-01-08 | 3 | 2.5 | 5392.5186 | 5202.0174 | Standard Error | 34.7036
2020-01-08 | 3 | 5 | 10616.6939 | 10426.1927 | Observations | 5
2019-11-20 | 1 | 0 | 68.7502 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.25 | 467.0077 | 398.2575 | Multiple R | 0.99997 | Regression | 1 | 34192760.6182 | 34192760.6182 | 44590.9507 | 2.3419e-07 | Intercept | 24.1563 | 18.3573 | 41.0481 | 1.3159 | 0.27973 | -34.2647 | 82.5773 | -34.2647 | 82.5773 | 0.089943 | 0.27256
2019-11-20 | 1 | 0.5 | 824.0238 | 755.2736 | R Square | 0.99993 | Residual | 3 | 2300.4282 | 766.8094 |  |  | X Variable 1 | 1506.0466 | 7.1321 |  | 211.1657 | 2.3419e-07 | 1483.3492 | 1528.744 | 1483.3492 | 1528.744
2019-11-20 | 1 | 1.25 | 2015.1422 | 1946.392 | Adjusted R Square | 0.99991 | Total | 4 | 34195061.0464
2019-11-20 | 1 | 2.5 | 3842.5174 | 3773.7672 | Standard Error | 27.6913
2019-11-20 | 1 | 5 | 7623.2842 | 7554.5341 | Observations | 5
2020-02-27 | 1 | 0 | 283.7528 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.25 | 889.5277 | 605.7749 | Multiple R | 0.99981 | Regression | 1 | 114363342.8466 | 114363342.8466 | 7875.752 | 3.1538e-06 | Intercept | -103.3891 | 79.8843 | 178.6268 | -1.2942 | 0.28621 | -357.6167 | 150.8385 | -357.6167 | 150.8385 | 0.21402 | 0.64853
2020-02-27 | 1 | 0.5 | 1583.3378 | 1299.5849 | R Square | 0.99962 | Residual | 3 | 43562.8279 | 14520.9426 |  |  | X Variable 1 | 2754.3223 | 31.0362 |  | 88.7454 | 3.1538e-06 | 2655.5512 | 2853.0934 | 2655.5512 | 2853.0934
2020-02-27 | 1 | 1.25 | 3687.7262 | 3403.9734 | Adjusted R Square | 0.99949 | Total | 4 | 114406905.6745
2020-02-27 | 1 | 2.5 | 6883.9084 | 6600.1556 | Standard Error | 120.5029
2020-02-27 | 1 | 5 | 14023.3805 | 13739.6276 | Observations | 5
2020-02-27 | 2 | 0 | 300.0032 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.25 | 881.7772 | 581.7741 | Multiple R | 0.99995 | Regression | 1 | 85225481.0775 | 85225481.0775 | 29547.4066 | 4.3415e-07 | Intercept | 11.1396 | 35.6032 | 79.6112 | 0.31288 | 0.77486 | -102.1658 | 124.4449 | -102.1658 | 124.4449 | 0.11049 | 0.33483
2020-02-27 | 2 | 0.5 | 1454.5741 | 1154.5709 | R Square | 0.9999 | Residual | 3 | 8653.0925 | 2884.3642 |  |  | X Variable 1 | 2377.6946 | 13.8324 |  | 171.8936 | 4.3415e-07 | 2333.6738 | 2421.7153 | 2333.6738 | 2421.7153
2020-02-27 | 2 | 1.25 | 3355.3941 | 3055.3909 | Adjusted R Square | 0.99986 | Total | 4 | 85234134.17
2020-02-27 | 2 | 2.5 | 6274.1277 | 5974.1246 | Standard Error | 53.7063
2020-02-27 | 2 | 5 | 12177.9392 | 11877.9361 | Observations | 5

## arsenic_As
Date | No_cal | Std_conc (ppb) | AsO (cps) | AsO (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 13.75 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.25 | 1647.5951 | 1633.845 | Multiple R | 0.99996 | Regression | 1 | 715167322.6116 | 715167322.6116 | 35330.4884 | 3.3205e-07 | Intercept | -179.9164 | 94.3177 | 210.9008 | -1.9076 | 0.15249 | -480.0774 | 120.2447 | -480.0774 | 120.2447 | 0.10105 | 0.3062
2020-01-08 | 1 | 0.5 | 3357.3946 | 3343.6446 | R Square | 0.99992 | Residual | 3 | 60726.6434 | 20242.2145 |  |  | X Variable 1 | 6887.715 | 36.6438 |  | 187.9641 | 3.3205e-07 | 6771.0982 | 7004.3319 | 6771.0982 | 7004.3319
2020-01-08 | 1 | 1.25 | 8339.434 | 8325.684 | Adjusted R Square | 0.99989 | Total | 4 | 715228049.2549
2020-01-08 | 1 | 2.5 | 16891.238 | 16877.488 | Standard Error | 142.2751
2020-01-08 | 1 | 5 | 34366.7993 | 34353.0493 | Observations | 5
2020-01-08 | 2 | 0 | 22.75 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.25 | 1463.825 | 1441.075 | Multiple R | 0.99994 | Regression | 1 | 546994195.3208 | 546994195.3208 | 26718.7398 | 5.0488e-07 | Intercept | -144.503 | 94.8523 | 212.0961 | -1.5235 | 0.22503 | -446.3652 | 157.3592 | -446.3652 | 157.3592 | 0.11619 | 0.3521
2020-01-08 | 2 | 0.5 | 2908.046 | 2885.296 | R Square | 0.99989 | Residual | 3 | 61416.9155 | 20472.3052 |  |  | X Variable 1 | 6023.6912 | 36.8515 |  | 163.4587 | 5.0488e-07 | 5906.4134 | 6140.969 | 5906.4134 | 6140.969
2020-01-08 | 2 | 1.25 | 7427.1803 | 7404.4303 | Adjusted R Square | 0.99985 | Total | 4 | 547055612.2363
2020-01-08 | 2 | 2.5 | 14724.8353 | 14702.0853 | Standard Error | 143.0815
2020-01-08 | 2 | 5 | 30092.4149 | 30069.6649 | Observations | 5
2020-01-08 | 3 | 0 | 27.75 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.25 | 1375.0662 | 1347.3162 | Multiple R | 0.99992 | Regression | 1 | 445811122.2733 | 445811122.2733 | 19730.5917 | 7.9558e-07 | Intercept | 16.0039 | 99.6482 | 222.8203 | 0.1606 | 0.88261 | -301.1213 | 333.1291 | -301.1213 | 333.1291 | 0.13521 | 0.40974
2020-01-08 | 3 | 0.5 | 2748.7645 | 2721.0145 | R Square | 0.99985 | Residual | 3 | 67784.7572 | 22594.9191 |  |  | X Variable 1 | 5438.095 | 38.7148 |  | 140.4656 | 7.9558e-07 | 5314.8873 | 5561.3027 | 5314.8873 | 5561.3027
2020-01-08 | 3 | 1.25 | 7016.2227 | 6988.4727 | Adjusted R Square | 0.9998 | Total | 4 | 445878907.0304
2020-01-08 | 3 | 2.5 | 13455.5844 | 13427.8344 | Standard Error | 150.3161
2020-01-08 | 3 | 5 | 27285.0341 | 27257.2841 | Observations | 5
2019-11-20 | 1 | 0 | 8.25 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.25 | 1344.0633 | 1335.8133 | Multiple R | 0.99991 | Regression | 1 | 423513033.8245 | 423513033.8245 | 17471.8806 | 9.5471e-07 | Intercept | -37.0285 | 103.2114 | 230.7878 | -0.35876 | 0.74353 | -365.4934 | 291.4364 | -365.4934 | 291.4364 | 0.14369 | 0.43542
2019-11-20 | 1 | 0.5 | 2641.7443 | 2633.4943 | R Square | 0.99983 | Residual | 3 | 72719.0812 | 24239.6937 |  |  | X Variable 1 | 5300.3522 | 40.0991 |  | 132.1812 | 9.5471e-07 | 5172.7389 | 5427.9655 | 5172.7389 | 5427.9655
2019-11-20 | 1 | 1.25 | 6669.0574 | 6660.8074 | Adjusted R Square | 0.99977 | Total | 4 | 423585752.9057
2019-11-20 | 1 | 2.5 | 12986.1525 | 12977.9025 | Standard Error | 155.691
2019-11-20 | 1 | 5 | 26568.4361 | 26560.1861 | Observations | 5
2020-02-27 | 1 | 0 | 11.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 1 | 0.25 | 2591.2351 | 2579.7351 | Multiple R | 0.9999 | Regression | 1 | 1779579086.5201 | 1779579086.5201 | 14495.0063 | 1.2634e-06 | Intercept | -306.0514 | 232.281 | 519.3961 | -1.3176 | 0.27923 | -1045.2733 | 433.1705 | -1045.2733 | 433.1705 | 0.15775 | 0.47804
2020-02-27 | 1 | 0.5 | 5132.422 | 5120.922 | R Square | 0.99979 | Residual | 3 | 368315.6223 | 122771.8741 |  |  | X Variable 1 | 10865.0064 | 90.2445 |  | 120.3952 | 1.2634e-06 | 10577.8081 | 11152.2047 | 10577.8081 | 11152.2047
2020-02-27 | 1 | 1.25 | 13426.5595 | 13415.0595 | Adjusted R Square | 0.99972 | Total | 4 | 1779947402.1424
2020-02-27 | 1 | 2.5 | 26346.0226 | 26334.5226 | Standard Error | 350.3882
2020-02-27 | 1 | 5 | 54248.5646 | 54237.0646 | Observations | 5
2020-02-27 | 2 | 0 | 12.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-02-27 | 2 | 0.25 | 2402.202 | 2389.702 | Multiple R | 0.99998 | Regression | 1 | 1455374571.4286 | 1455374571.4286 | 65668.3659 | 1.3104e-07 | Intercept | 9.4575 | 98.6901 | 220.6778 | 0.09583 | 0.9297 | -304.6184 | 323.5334 | -304.6184 | 323.5334 | 0.074116 | 0.22459
2020-02-27 | 2 | 0.5 | 4843.0709 | 4830.5709 | R Square | 0.99995 | Residual | 3 | 66487.4732 | 22162.4911 |  |  | X Variable 1 | 9825.5921 | 38.3425 |  | 256.2584 | 1.3104e-07 | 9703.5691 | 9947.6151 | 9703.5691 | 9947.6151
2020-02-27 | 2 | 1.25 | 12528.9918 | 12516.4918 | Adjusted R Square | 0.99994 | Total | 4 | 1455441058.9018
2020-02-27 | 2 | 2.5 | 24558.0918 | 24545.5918 | Standard Error | 148.8707
2020-02-27 | 2 | 5 | 49120.5558 | 49108.0558 | Observations | 5

## cadmium_Cd
Date | No_cal | Std_conc (ppb) | Cd (cps) | Cd (BLK corrected) | SUMMARY OUTPUT |  | ANOVA
2020-01-08 | 1 | 0 | 28 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | LOD (µg/L) | LOQ (µg/L)
2020-01-08 | 1 | 0.01 | 172 | 144 | Multiple R | 0.99986 | Regression | 1 | 4915953.1648 | 4915953.1648 | 11025.3229 | 1.9043e-06 | Intercept | -13.7977 | 13.9982 | 31.3009 | -0.98567 | 0.39697 | -58.3462 | 30.7509 | -58.3462 | 30.7509 | 0.0072353 | 0.021925
2020-01-08 | 1 | 0.02 | 311 | 283 | R Square | 0.99973 | Residual | 3 | 1337.6352 | 445.8784 |  |  | X Variable 1 | 14276.2852 | 135.9626 |  | 105.0015 | 1.9043e-06 | 13843.5915 | 14708.979 | 13843.5915 | 14708.979
2020-01-08 | 1 | 0.05 | 710 | 682 | Adjusted R Square | 0.99964 | Total | 4 | 4917290.8
2020-01-08 | 1 | 0.1 | 1420 | 1392 | Standard Error | 21.1158
2020-01-08 | 1 | 0.2 | 2883 | 2855 | Observations | 5
2020-01-08 | 2 | 0 | 24 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 2 | 0.01 | 140 | 116 | Multiple R | 0.99999 | Regression | 1 | 3445304.5599 | 3445304.5599 | 116605.3498 | 5.5383e-08 | Intercept | -2.7197 | 3.6035 | 8.0576 | -0.75476 | 0.50524 | -14.1875 | 8.7481 | -14.1875 | 8.7481 | 0.0022248 | 0.0067418
2020-01-08 | 2 | 0.02 | 261 | 237 | R Square | 0.99997 | Residual | 3 | 88.6401 | 29.5467 |  |  | X Variable 1 | 11951.5755 | 34.9998 |  | 341.4753 | 5.5383e-08 | 11840.1904 | 12062.9605 | 11840.1904 | 12062.9605
2020-01-08 | 2 | 0.05 | 624 | 600 | Adjusted R Square | 0.99997 | Total | 4 | 3445393.2
2020-01-08 | 2 | 0.1 | 1209 | 1185 | Standard Error | 5.4357
2020-01-08 | 2 | 0.2 | 2414 | 2390 | Observations | 5
2020-01-08 | 3 | 0 | 23 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2020-01-08 | 3 | 0.01 | 118 | 95 | Multiple R | 0.99995 | Regression | 1 | 2722603.1343 | 2722603.1343 | 28081.036 | 4.6859e-07 | Intercept | -2.4527 | 6.5275 | 14.596 | -0.37575 | 0.7321 | -23.2263 | 18.3208 | -23.2263 | 18.3208 | 0.0045336 | 0.013738
2020-01-08 | 3 | 0.02 | 231 | 208 | R Square | 0.99989 | Residual | 3 | 290.8657 | 96.9552 |  |  | X Variable 1 | 10624.3781 | 63.4011 |  | 167.574 | 4.6859e-07 | 10422.6074 | 10826.1488 | 10422.6074 | 10826.1488
2020-01-08 | 3 | 0.05 | 566 | 543 | Adjusted R Square | 0.99986 | Total | 4 | 2722894
2020-01-08 | 3 | 0.1 | 1082 | 1059 | Standard Error | 9.8466
2020-01-08 | 3 | 0.2 | 2143 | 2120 | Observations | 5
2019-11-20 | 1 | 0 | 2.5 | 0 | Regression Statistics |  |  | df | SS | MS | F | Significance F |  | Coefficients | SE | SD | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
2019-11-20 | 1 | 0.01 | 85.5003 | 83.0003 | Multiple R | 0.99986 | Regression | 1 | 1364489.9978 | 1364489.9978 | 10797.5436 | 1.9649e-06 | Intercept | 15.1482 | 7.4522 | 16.6637 | 2.0327 | 0.13499 | -8.5682 | 38.8645 | -8.5682 | 38.8645 | 0.0073112 | 0.022155
2019-11-20 | 1 | 0.02 | 164.501 | 162.001 | R Square | 0.99972 | Residual | 3 | 379.1112 | 126.3704 |  |  | X Variable 1 | 7521.3626 | 72.3826 |  | 103.9112 | 1.9649e-06 | 7291.0089 | 7751.7163 | 7291.0089 | 7751.7163
2019-11-20 | 1 | 0.05 | 397.0055 | 394.5055 | Adjusted R Square | 0.99963 | Total | 4 | 1364869.1091
2019-11-20 | 1 | 0.1 | 785.2716 | 782.7716 | Standard Error | 11.2415
2019-11-20 | 1 | 0.2 | 1514.0803 | 1511.5803 | Observations | 5

## Distribution_sample_std
Element concentrations (unfactored)
Sample ID | Co 59 (ug/L) | Ni 60 (ug/L) | Zn 66 (ug/L) | Cu 65 (ug/L) | Sr 88 (ug/L) | Mo 98 (ug/L) | Sb 121 (ug/L) | Cs 133 (ug/L) | Ba 138 (ug/L) | W 184 (ug/L) | Tl 205 (ug/L) | Pb 208 (ug/L) | U 238 (ug/L) | Mn 55 (ug/L) | Se 78 (ug/L) | AsO 91 (ug/L) | Cd 111 (ug/L) | Cd-1 112 (ug/L) | Cd-1 113 (ug/L) | Cd 114 (ug/L)
7014562 | 0.0092998 | 0.02343 | 64.4693 | 0.24445 | 1.7415 | 1.4578 | 0.0028453 | 0.10569 | 0.058354 | 0.0025441 | 0.0032052 | 0.034993 | 0.00014525 | 0.012614 | 1.269 | 0.40896 | 0.0077772 | 0.0056576 | 0.0069609 | 0.0051125
5018625 | 0.0066464 | 0.021478 | 78.7637 | 0.32334 | 2.8524 | 0.94099 | 4.3596e-05 | 0.042367 | 0.040464 | 0.0011331 | 0.0041331 | 0.041969 | 0.00020253 | 0.011816 | 0.62386 | 0.065303 | 0.029279 | 0.027889 | 0.029102 | 0.027784
5018625 | 0.0063448 | 0.014734 | 78.3482 | 0.31538 | 2.8372 | 0.93432 | -1.9809e-05 | 0.043973 | 0.039971 | 0.0013021 | 0.0038597 | 0.0419 | 0.00015592 | 0.0124 | 0.629 | 0.061688 | 0.028234 | 0.027702 | 0.028243 | 0.026531
6014755 | 0.011448 | 0.05732 | 21.1403 | 0.32875 | 5.2122 | 2.0421 | 0.0014094 | 0.20649 | 0.055321 | 0.0050731 | 0.0058727 | 0.036442 | 0.00013178 | 0.0049618 | 1.3135 | 0.38997 | 0.013879 | 0.012057 | 0.014636 | 0.011422
6016405 | 0.010477 | 0.034508 | 26.4799 | 0.33937 | 2.9785 | 0.95622 | 6.0341e-05 | 0.1135 | 0.05124 | 0.0054298 | 0.0047046 | 0.015125 | 8.9363e-05 | 0.0078302 | 1.5918 | 0.29835 | 0.011257 | 0.0090241 | 0.010366 | 0.0091537
7017480 | 0.022168 | 1.4219 | 0.16362 | 1.9221 | 1.9221 | 0.39844 | -0.00026756 | 0.076588 | 0.04462 | 0.0025264 | 0.0027214 | 0.0057913 | 8.3576e-05 | 0.0070699 | 0.44555 | 0.055271 | 0.0073392 | 0.0069289 | 0.0089033 | 0.0062062
7017480 | 0.024066 | 1.5919 | 0.16824 | 1.8496 | 1.8496 | 0.39803 | -0.00055359 | 0.076382 | 0.042343 | 0.0019577 | 0.0027835 | 0.0076836 | 5.8173e-05 | 0.0051621 | 0.4586 | 0.055456 | 0.007641 | 0.0065913 | 0.0074722 | 0.0062719
7017979 | 0.015425 | 17.7548 | 0.17521 | 1.2781 | 1.2781 | 0.44893 | -0.00047338 | 0.096135 | 0.013773 | 0.00058127 | 0.0022078 | 0.013227 | 7.1528e-05 | 0.0077004 | 1.1882 | 0.51144 | 0.0068494 | 0.0059969 | 0.0075284 | 0.0053537
7017979 | 0.014893 | 18.0288 | 0.14044 | 1.2632 | 1.2632 | 0.44156 | -0.0011699 | 0.094443 | 0.011855 | 0.00036067 | 0.0020149 | 0.0115 | 4.261e-05 | 0.0062189 | 1.1973 | 0.50688 | 0.0080283 | 0.0059265 | 0.0073357 | 0.0055891
3010112 | 0.01254 | 8.4827 | 0.15211 | 0.64487 | 0.64487 | 0.51341 | 2.8464e-05 | 0.028357 | 0.021109 | 0.00063362 | 0.0066867 | 0.0065035 | 3.4953e-05 | 0.0047878 | 0.30665 | 0.14161 | 0.003712 | 0.0028957 | 0.0036897 | 0.0026123
3010112 | 0.011175 | 17.6369 | 0.14421 | 0.63767 | 0.63767 | 0.50785 | 3.5744e-05 | 0.027679 | 0.020015 | 0.00079949 | 0.0066665 | 0.0088782 | 3.5864e-05 | 0.0044178 | 0.28797 | 0.14158 | 0.0029989 | 0.0027107 | 0.0027008 | 0.0024583
3014819 | 0.02176 | 14.1116 | 0.16064 | 0.38864 | 0.38864 | 0.54592 | -0.00040959 | 0.073594 | 0.0052517 | 0.0013219 | 0.0027536 | 0.013911 | 5.9008e-05 | 0.0054454 | 0.5494 | 0.2059 | 0.007551 | 0.007629 | 0.0075854 | 0.0069382
3014819 | 0.021703 | 15.1619 | 0.15491 | 0.39132 | 0.39132 | 0.55092 | -0.00028607 | 0.075068 | 0.0052394 | 0.0013778 | 0.002688 | 0.014356 | 6.5588e-05 | 0.0055429 | 0.53618 | 0.20075 | 0.0074606 | 0.0085295 | 0.0080281 | 0.0070154
3019179 | 0.024835 | 26.7827 | 0.22623 | 2.4485 | 2.4485 | 0.73104 | 0.00043399 | 0.06448 | 0.040191 | 0.0009589 | 0.002984 | 0.025183 | 3.9329e-05 | 0.0066847 | 0.75753 | 0.19712 | 0.02156 | 0.020424 | 0.021833 | 0.020267
3019179 | 0.02789 | 52.7507 | 0.20943 | 2.4154 | 2.4154 | 0.73026 | 8.2902e-05 | 0.063641 | 0.039355 | 0.000921 | 0.003007 | 0.027374 | 3.6051e-05 | 0.0089915 | 0.72378 | 0.19285 | 0.022343 | 0.019067 | 0.020196 | 0.02114
5011752 | 0.03001 | 6.3424 | 0.13453 | 1.391 | 1.391 | 0.2808 | 0.0029033 | 0.045251 | 0.0091253 | 0.00018649 | 0.0017744 | 0.007629 | 6.7182e-06 | 0.003447 | 0.53057 | 0.23276 | 0.0083173 | 0.0073418 | 0.0083805 | 0.0073046
5011752 | 0.029196 | 6.3226 | 0.12151 | 1.3806 | 1.3806 | 0.27825 | 0.0030628 | 0.044802 | 0.0087469 | 0.0001299 | 0.0018762 | 0.0073631 | 2.2729e-05 | 0.002996 | 0.50509 | 0.23992 | 0.0080155 | 0.0071801 | 0.0079314 | 0.0067424
6013040 | 0.01217 | 0.030216 | 19.4974 | 0.45678 | 2.2824 | 1.2926 | 0.0012749 | 0.15371 | 0.035504 | 0.0028052 | 0.0030925 | 0.047617 | 8.8334e-05 | 0.014888 | 2.2834 | 0.23316 | 0.01806 | 0.016442 | 0.018179 | 0.015039
6013040 | 0.01127 | 0.032582 | 19.6725 | 0.46082 | 2.3009 | 1.3118 | 0.00080032 | 0.15697 | 0.034679 | 0.0027421 | 0.0030159 | 0.047394 | 0.00010164 | 0.0070131 | 2.3194 | 0.25206 | 0.019463 | 0.016861 | 0.018697 | 0.015449
6022995 | 0.0081574 | 0.033578 | 13.8709 | 0.29347 | 2.5932 | 0.8322 | -0.00054505 | 0.09999 | 0.030932 | 0.001087 | 0.0019398 | 0.032761 | 4.2009e-05 | 0.0046106 | 1.0967 | 0.078116 | 0.0045944 | 0.003527 | 0.0041676 | 0.003176
6022995 | 0.0080034 | 0.030369 | 12.7349 | 0.27046 | 2.3966 | 0.78247 | -0.00061532 | 0.095857 | 0.024519 | 0.00072907 | 0.0018592 | 0.025021 | 9.3952e-05 | 0.0028196 | 1.0084 | 0.073469 | 0.0041498 | 0.002795 | 0.0035861 | 0.002834
8022461 | 0.0051961 | 0.0147 | 40.1942 | 0.17946 | 0.82404 | 0.7515 | -0.0010806 | 0.10919 | 2.4902 | 0.0014218 | 0.0019305 | 0.026486 | 9.309e-05 | 0.0090602 | 0.6495 | 0.15859 | 0.0066824 | 0.0050052 | 0.0066064 | 0.0050519
8022461 | 0.0055996 | 0.018017 | 56.9343 | 0.17781 | 0.85283 | 0.75265 | -0.00040559 | 0.1105 | 2.2512 | 0.0012856 | 0.0019593 | 0.033162 | 0.00015801 | 0.015735 | 0.64629 | 0.15525 | 0.0070457 | 0.0062201 | 0.0068362 | 0.0055638
5021103 | 0.0023063 | 0.017422 | 2.6918 | 0.084359 | 0.21559 | 0.16997 | -0.0018869 | 0.03313 | 0.0060191 | 0.00013784 | 0.00025409 | 0.0047506 | 1.2063e-05 | 0.0014586 | 0.19686 | 0.043302 | 0.0060084 | 0.0060667 | 0.007132 | 0.0062116
5021103 | 0.002366 | 0.014632 | 2.5734 | 0.082838 | 0.2113 | 0.16363 | -0.0014153 | 0.03286 | 0.005147 | 0.00022194 | 0.00021999 | 0.0053526 | 1.9464e-05 | 0.0013053 | 0.18667 | 0.039932 | 0.0056542 | 0.0064229 | 0.0058673 | 0.0055155
6022863 | 0.0067316 | 0.087343 | 7.8903 | 0.25382 | 1.1851 | 1.4802 | 0.0021213 | 0.1276 | 0.009947 | 0.0021688 | 0.0031563 | 0.0497 | 0.00052727 | 0.015238 | 1.0073 | 0.64124 | 0.010567 | 0.0079933 | 0.0093122 | 0.0015325
6022863 | 0.0070216 | 0.078999 | 7.3174 | 0.25051 | 1.1669 | 1.4489 | 0.0024036 | 0.12503 | 0.0095615 | 0.0021426 | 0.0031901 | 0.047879 | 0.0004832 | 0.011123 | 0.96491 | 0.62812 | 0.010951 | 0.0073037 | 0.010159 | 0.0017436
7016034 | 0.0043857 | 0.0053311 | 4.9319 | 0.086539 | 1.189 | 0.38049 | -0.00073063 | 0.030191 | 0.11878 | 0.00032568 | 0.00057817 | 0.0091888 | 5.8916e-05 | 0.0035501 | 0.46114 | 0.029024 | 0.0028277 | 0.0021595 | 0.0035956 | 0.0020524
7016034 | 0.0037345 | 0.002323 | 4.7414 | 0.078602 | 1.1487 | 0.36852 | -0.00051263 | 0.029265 | 0.033024 | 0.00027427 | 0.00058943 | 0.0084228 | 5.7536e-05 | 0.0035457 | 0.42936 | 0.026955 | 0.002124 | 0.0015133 | 0.0033176 | 0.0021417
7013655 | 0.0056656 | 0.013215 | 79.9915 | 0.36938 | 0.82342 | 0.18987 | -0.00078899 | 0.11118 | 0.013035 | 0.0010951 | 0.002674 | 0.023427 | 6.3874e-05 | 0.015453 | 0.74044 | 0.081239 | 0.0057044 | 0.0054431 | 0.0056661 | 0.0044032
7013655 | 0.0053755 | 0.013988 | 85.926 | 0.36692 | 0.79839 | 0.18883 | -0.00093964 | 0.10855 | 0.013118 | 0.0014199 | 0.0025486 | 0.024388 | 5.3173e-05 | 0.016291 | 0.75829 | 0.083496 | 0.0067706 | 0.004258 | 0.0058342 | 0.0046433
7016050 | 0.007224 | 0.041785 | 14.5991 | 0.28226 | 3.789 | 1.4275 | 0.0013552 | 0.13829 | 0.025306 | 0.0013452 | 0.0020198 | 0.061879 | 0.00010631 | 0.0031757 | 1.3089 | 0.92049 | 0.029225 | 0.030408 | 0.029331 | 0.029368
7016050 | 0.0067044 | 0.040986 | 14.3136 | 0.27771 | 3.6786 | 1.3847 | 0.001829 | 0.13866 | 0.021842 | 0.0015329 | 0.001969 | 0.060508 | 0.000101 | 0.0035858 | 1.2685 | 0.87757 | 0.029779 | 0.029366 | 0.030364 | 0.027443
3015106 | 0.0050977 | 0.063558 | 6.351 | 0.2771 | 0.53547 | 0.89296 | 0.00048059 | 0.10874 | 0.0080946 | 0.0013939 | 0.0030301 | 0.017852 | 5.3983e-05 | 0.0030048 | 1.4077 | 0.97508 | 0.0079526 | 0.0044551 | 0.0086313 | 0.0046676
3015106 | 0.0046923 | 0.060721 | 6.0346 | 0.25357 | 0.50655 | 0.85026 | 0.0010556 | 0.1049 | 0.0075979 | 0.001368 | 0.0030743 | 0.016699 | 6.6844e-05 | 0.0043757 | 1.3583 | 0.93857 | 0.0083681 | 0.0057074 | 0.0080779 | 0.0042711
3017494 | 0.0092719 | 0.051166 | 13.7059 | 0.27708 | 1.9667 | 0.39118 | 0.000445 | 0.091914 | 0.033113 | 0.00042863 | 0.0021277 | 0.014253 | 5.7213e-05 | 0.0056354 | 0.89884 | 0.73365 | 0.0057632 | 0.0039975 | 0.0050679 | 0.0042158
3017494 | 0.0091721 | 0.052154 | 13.919 | 0.23229 | 1.9526 | 0.39052 | -0.00070244 | 0.089705 | 0.033698 | 0.00048293 | 0.002411 | 0.012253 | 9.7517e-05 | 0.0041288 | 0.87247 | 0.73812 | 0.0056664 | 0.004608 | 0.0056211 | 0.0039101
3014967 | 0.0063693 | 0.030031 | 14.5822 | 0.22306 | 0.89929 | 3.3661 | 0.0011311 | 0.068495 | 0.019803 | 0.0021731 | 0.0020136 | 0.018547 | 8.9309e-05 | 0.0033715 | 0.9177 | 1.0537 | 0.0059073 | 0.0046026 | 0.0062314 | 0.004286
3014967 | 0.0061084 | 0.02626 | 14.076 | 0.19638 | 0.8562 | 3.2259 | 0.00051504 | 0.068441 | 0.0067234 | 0.0022962 | 0.0020579 | 0.017074 | 5.6912e-05 | 0.0032966 | 0.85499 | 1.0034 | 0.0067431 | 0.0041969 | 0.0061665 | 0.0042153
3016170 | 0.008615 | 0.033139 | 12.0886 | 0.25607 | 1.8069 | 0.60783 | -0.0003767 | 0.095393 | 0.01612 | 0.0089406 | 0.00596 | 0.027296 | 3.4636e-05 | 0.010945 | 1.0151 | 0.84396 | 0.0082254 | 0.0079269 | 0.0086712 | 0.0054403
3016170 | 0.0089536 | 0.030824 | 12.1692 | 0.23825 | 1.7143 | 0.5717 | -0.00065467 | 0.091364 | 0.014033 | 0.0091839 | 0.0055746 | 0.026105 | 4.1828e-05 | 0.010452 | 0.93822 | 0.81424 | 0.0077135 | 0.0069569 | 0.0085182 | 0.0060815
3021530 | 5.0702 | 6.9125 | 31.2699 | 11.4263 | 1.4644 | 8.6043 | 1.7602 | 1.2266 | -0.0045037 | 1.0684 | 0.78287 | 10.4321 | -9.4798e-05 | 0.77648 | 3366.1344 | 27.1214 | 4.3291 | 0.37614 | 0.97787 | 0.57876
3021530 | 0.0030173 | 0.01151 | 4.3847 | 0.13078 | 0.44816 | 0.34424 | -0.00077207 | 0.056307 | 0.0064423 | 0.00054136 | 0.0010474 | 0.0058746 | -7.6849e-05 | 0.0088427 | 0.27309 | 0.043489 | 0.0026302 | 0.0025481 | 0.0027542 | 0.0014664
4010272 | 0.0068097 | 0.21972 | 24.1679 | 0.30386 | 3.1515 | 0.16818 | -0.00036871 | 0.060674 | 0.059568 | 0.00091489 | 0.0018575 | 0.014382 | -1.9269e-05 | 0.015728 | 0.43495 | 0.24429 | 0.0077439 | 0.0081718 | 0.0082344 | 0.0064967
4010272 | 0.0061824 | 0.0035879 | 26.1957 | 0.28007 | 3.1229 | 0.16212 | -0.00047661 | 0.061935 | 0.05452 | 0.0006268 | 0.0019226 | 0.014683 | -4.8025e-05 | 0.0080522 | 0.41702 | 0.25388 | 0.0082734 | 0.016822 | 0.0079133 | 0.0058609
5012473 | 0.0072613 | 0.04543 | 13.6462 | 0.30092 | 1.6173 | 0.45023 | -0.00032576 | 0.1762 | 0.029169 | 0.00023421 | 0.0037065 | 0.027906 | 0.00015802 | 0.0099513 | 0.69622 | 1.0015 | 0.028146 | 0.027279 | 0.02934 | 0.026671
5012473 | 0.0074594 | 0.043415 | 13.7251 | 0.2975 | 1.633 | 0.44735 | 0.00065738 | 0.17911 | 0.013594 | 0.00022541 | 0.0039257 | 0.028092 | 0.00015851 | 0.0034931 | 0.7008 | 1.0145 | 0.0292 | 0.027297 | 0.028113 | 0.027713
5019737 | 0.0017775 | 0.0089337 | 2.3338 | 0.072098 | 0.44415 | 0.25287 | -0.00076229 | 0.018374 | 0.0046515 | 0.001814 | 6.3457e-05 | 0.0051913 | 2.882e-05 | 0.0027815 | 0.27533 | 0.049938 | 0.0067875 | 0.0067503 | 0.0060481 | 0.0054186
5019737 | 0.0019018 | 0.0091996 | 2.2918 | 0.066662 | 0.43295 | 0.24561 | -0.00089822 | 0.017386 | 0.0052524 | 0.0019628 | 0.00010339 | 0.0064385 | -3.4665e-05 | 0.0036009 | 0.2876 | 0.049497 | 0.0062236 | 0.0057737 | 0.0060507 | 0.0048861
5022428 | 0.0046961 | 0.016448 | 6.2456 | 0.16876 | 0.94185 | 0.19542 | -0.00050596 | 0.10285 | 0.0065095 | 0.0012773 | 0.0026362 | 0.029636 | -3.6691e-05 | 0.0033628 | 0.58665 | 0.52825 | 0.0067057 | 0.0050877 | 0.0066762 | 0.0052426
5022428 | 0.0044443 | 0.014507 | 6.359 | 0.16757 | 0.92851 | 0.1882 | -0.0007299 | 0.10049 | 0.0061209 | 0.0015093 | 0.0025684 | 0.029742 | -4.8242e-05 | 0.0035802 | 0.58597 | 0.52601 | 0.006869 | 0.0051029 | 0.0068687 | 0.0048236
6012213 | 0.01672 | 0.028192 | 18.1139 | 0.26089 | 2.1142 | 1.2299 | 0.007714 | 0.14424 | 0.015557 | 0.0044597 | 0.0025681 | 0.015657 | 0.00081297 | 0.0073273 | 1.3045 | 0.10988 | 0.012492 | 0.0086606 | 0.01043 | 0.008554
6012213 | 0.016164 | 0.011152 | 17.6139 | 0.24009 | 2.103 | 1.2049 | -0.00042463 | 0.13971 | 0.018306 | 0.0013265 | 0.0014365 | 0.012604 | -1.7031e-05 | 0.020851 | 1.2964 | 0.11137 | 0.011728 | 0.009626 | 0.010838 | 0.0092104
6017142 | 0.012322 | 0.024189 | 12.2703 | 0.18619 | 1.5171 | 0.71775 | -0.00013897 | 0.084038 | 0.010532 | 0.0019203 | 0.0016008 | 0.0075479 | -6.397e-05 | 0.0077359 | 0.7137 | 0.15495 | 0.006784 | 0.0057076 | 0.0075994 | 0.0048728
6017142 | 0.011861 | 0.024713 | 13.6395 | 0.22299 | 1.5256 | 0.72808 | -0.00011092 | 0.084144 | 0.0092671 | 0.0018043 | 0.0016762 | 0.0089705 | -5.2132e-05 | 0.0048052 | 0.69248 | 0.16157 | 0.0075307 | 0.0051952 | 0.0066793 | 0.0046892
6022103 | 0.019194 | 0.23154 | 11.057 | 0.33753 | 2.3704 | 1.491 | 0.0010745 | 0.19791 | 0.027923 | 0.02192 | 0.0034942 | 0.045644 | 0.00010676 | 0.0049428 | 2.1718 | 2.4339 | 0.013802 | 0.0097084 | 0.012721 | 0.0064679
6022103 | 0.018637 | 0.22355 | 10.9359 | 0.33463 | 2.3767 | 1.4861 | 0.0007411 | 0.19859 | 0.02731 | 0.021696 | 0.0035537 | 0.045033 | 7.8377e-05 | 0.0063676 | 2.1205 | 2.4418 | 0.012848 | 0.0093512 | 0.01211 | 0.0054935
4016912 | 0.018166 | 0.054285 | 14.0569 | 0.36281 | 2.5063 | 1.2319 | 0.00025776 | 0.20351 | 0.033543 | 0.0023895 | 0.0080933 | 0.028241 | 0.00017989 | 0.008802 | 1.2686 | 0.54669 | 0.0096133 | 0.0069853 | 0.0099529 | 0.0060842
4016912 | 0.018661 | 0.049757 | 14.2635 | 0.34011 | 2.4416 | 1.2062 | 0.00024333 | 0.19862 | 0.032864 | 0.0024911 | 0.0078442 | 0.027402 | 0.00014606 | 0.0072117 | 1.2453 | 0.5166 | 0.009055 | 0.0071144 | 0.0091138 | 0.0063313
4016955 | 0.0045904 | 0.042547 | 15.1405 | 0.26901 | 0.44796 | 1.2196 | 0.00041539 | 0.16145 | 0.0074975 | 0.0013146 | 0.0045358 | 0.018875 | -3.4546e-05 | 0.0073516 | 1.1949 | 1.0248 | 0.017539 | 0.01534 | 0.017567 | 0.014164
4016955 | 0.0048926 | 0.044295 | 14.6444 | 0.25718 | 0.425 | 1.1959 | 0.00057486 | 0.15737 | 0.0065366 | 0.0011171 | 0.0041634 | 0.018775 | 1.0553e-05 | 0.0077162 | 1.1619 | 1.0035 | 0.017505 | 0.015216 | 0.016874 | 0.013979
4022742 | 0.011009 | 0.0020161 | 27.2397 | 0.36433 | 4.2263 | 0.62824 | 0.0010871 | 0.17578 | 0.046566 | 0.0015248 | 0.0031733 | 0.059103 | 0.00013651 | 0.0093659 | 0.78775 | 0.15056 | 0.012488 | 0.0088917 | 0.012436 | 0.0019194
4022742 | 0.010258 | 5.3053e-05 | 30.8432 | 0.31227 | 4.0819 | 0.61315 | 0.00055928 | 0.16956 | 0.044371 | 0.0018065 | 0.0032152 | 0.05638 | 9.9909e-05 | 0.0089382 | 0.75553 | 0.14736 | 0.011986 | 0.006909 | 0.01228 | 0.0021266
3021041 | 0.0068981 | -0.0017467 | 6.8924 | 0.25588 | 1.9396 | 0.6457 | 0.0013098 | 0.052393 | 0.045424 | 0.00095071 | 0.0014221 | 0.1357 | -5.1705e-06 | 0.0054833 | 1.1456 | 0.2398 | 0.014582 | 0.012524 | 0.015265 | 0.012707
3021041 | 0.0061784 | -0.0034384 | 6.8418 | 0.25085 | 1.9253 | 0.63665 | 0.001298 | 0.051551 | 0.038959 | 0.00090181 | 0.001372 | 0.13394 | 6.9729e-06 | 0.0048711 | 1.1397 | 0.23364 | 0.013548 | 0.013565 | 0.013943 | 0.01142
6015247 | 0.019128 | 0.021249 | 13.6108 | 0.44686 | 3.9395 | 1.7792 | -0.0026403 | 0.10342 | 0.2975 | 0.0018293 | 0.0033302 | 0.041625 | 0.00010675 | 0.010065 | 1.9884 | 0.095719 | 0.02547 | 0.020516 | 0.023466 | 0.0193
8019959 | 0.013385 | 0.015306 | 14.7007 | 0.36269 | 3.5403 | 1.5686 | -0.0036717 | 0.16819 | 0.050075 | 0.0016484 | 0.0032229 | 0.024177 | 0.00020584 | 0.0073663 | 1.7522 | 0.70025 | 0.012988 | 0.0071531 | 0.012688 | 0.0070865
8019959 | 0.014944 | 0.026149 | 14.7538 | 0.37977 | 3.4906 | 1.5686 | -0.0026988 | 0.1637 | 0.13443 | 0.0016345 | 0.0033425 | 0.024984 | 0.00019934 | 0.013595 | 1.7565 | 0.69748 | 0.012844 | 0.0065717 | 0.013173 | 0.0067767
7018509 | 0.005708 | 0.016408 | 20.9582 | 0.20616 | 0.87508 | 0.2897 | -0.0046766 | 0.035027 | 0.022185 | 0.00091704 | 0.0019271 | 0.019357 | 9.7358e-05 | 0.0031986 | 0.58035 | 0.14482 | 0.01425 | 0.013352 | 0.013641 | 0.012977
7018509 | 0.0057553 | 0.015202 | 20.3644 | 0.20526 | 0.88421 | 0.27986 | -0.0039019 | 0.034625 | 0.021206 | 0.00077238 | 0.0019682 | 0.019165 | 9.5555e-05 | 0.0043631 | 0.57468 | 0.1451 | 0.014708 | 0.012965 | 0.013487 | 0.013024
8018030 | 0.0053679 | 0.0058942 | 24.3421 | 0.054497 | 1.2473 | 0.32755 | -0.0039987 | 0.023951 | 0.014411 | 0.00066833 | 0.00073405 | 0.030157 | 9.3224e-05 | 0.0072482 | 0.36662 | 0.026314 | 0.005737 | 0.0087456 | 0.005838 | 0.0030701
8018030 | 0.0052451 | 0.019308 | 24.829 | 0.053756 | 1.2765 | 0.33341 | -0.0029177 | 0.024669 | 0.015063 | 0.00085022 | 0.00061988 | 0.031005 | 8.4935e-05 | 0.0047432 | 0.37609 | 0.029134 | 0.0055492 | 0.0029505 | 0.0058284 | 0.0033057
5022860 | 0.003495 | 0.027879 | 0.33789 | 0.075568 | 0.84578 | 0.40958 | -0.0047771 | 0.023007 | 0.01227 | 0.00034108 | 0.00073235 | 0.0060561 | 2.2593e-05 | 0.001312 | 0.30995 | 0.021145 | 0.0034875 | 0.001353 | 0.0035549 | 0.0015937
5022860 | 0.003641 | 0.016311 | 0.34568 | 0.098291 | 0.82351 | 0.39441 | -0.0034538 | 0.022894 | 0.035937 | 0.00013034 | 0.00074742 | 0.0049269 | 3.7235e-05 | 0.0032387 | 0.31329 | 0.022088 | 0.0040382 | 0.0027709 | 0.003762 | 0.0016516
3015327 | 0.024587 | 0.023273 | 16.2188 | 0.26593 | 5.3666 | 0.77972 | -0.0040794 | 0.20772 | 0.067017 | 0.0010458 | 0.0030922 | 0.036083 | 0.00024685 | 0.0079085 | 1.819 | 2.2301 | 0.011542 | 0.0067111 | 0.011518 | 0.0064885
3015327 | 0.023616 | 0.024004 | 16.3193 | 0.27362 | 5.4126 | 0.78374 | -0.0041802 | 0.20499 | 0.064771 | 0.0010326 | 0.0032777 | 0.036582 | 0.0002833 | 0.0057291 | 1.789 | 2.2558 | 0.011354 | 0.0064367 | 0.011274 | 0.0062206
5012953 | 0.016895 | 0.029256 | 8.6638 | 0.39093 | 2.5175 | 2.1694 | -0.0036775 | 0.18295 | 0.11571 | 0.0061412 | 0.0046168 | 0.075542 | 4.4428e-05 | 0.013172 | 3.2847 | 0.28696 | 0.022594 | 0.017725 | 0.021767 | 0.016148
5012953 | 0.016632 | 0.019491 | 8.4675 | 0.38521 | 2.5859 | 2.1957 | -0.003985 | 0.18403 | 0.1096 | 0.0063237 | 0.0047879 | 0.065872 | 6.0109e-05 | 0.0067477 | 3.2917 | 0.2859 | 0.021356 | 0.016255 | 0.02116 | 0.016335
5022622 | 0.011565 | 0.0020062 | 12.4488 | 0.22058 | 2.0062 | 0.97138 | -0.002602 | 0.058098 | 0.031071 | 0.0012891 | 0.0030957 | 0.01827 | 6.2133e-05 | 0.0040642 | 0.71924 | 0.23189 | 0.009705 | 0.0054261 | 0.009529 | 0.0049324
5022622 | 0.011598 | 0.010898 | 12.3849 | 0.24453 | 2.0587 | 0.96832 | -0.002845 | 0.060332 | 0.028174 | 0.0011917 | 0.0029962 | 0.018464 | 9.5453e-05 | 0.0049361 | 0.66796 | 0.23132 | 0.009765 | 0.0049048 | 0.0098299 | 0.0051975
6012663 | 0.005532 | 0.0042701 | 1.6811 | 0.091581 | 0.92533 | 0.25256 | -0.001802 | 0.038104 | 0.023784 | 0.00029489 | 0.0012359 | 0.012229 | 2.9962e-05 | 0.010309 | 0.48176 | 0.037339 | 0.0038117 | 0.0015749 | 0.0032032 | 0.0018775
6012663 | 0.0056203 | 0.00032496 | 1.4517 | 0.0818 | 0.92925 | 0.25757 | -0.0051078 | 0.038418 | 0.019262 | 0.00022248 | 0.0011579 | 0.0079061 | 3.746e-05 | 0.0049926 | 0.47772 | 0.038717 | 0.0025879 | 0.0017749 | 0.0027477 | 0.0014774
7016344 | 0.0073821 | 0.1099 | 39.3847 | 0.68349 | 0.20342 | 1.3858 | -0.001859 | 0.10784 | 0.011983 | 0.0024595 | 0.0017398 | 0.090426 | 0.00028037 | 0.0074275 | 1.1361 | 0.23946 | 0.040923 | 0.040708 | 0.037117 | 0.034943
7016344 | 0.0069562 | 0.1049 | 39.0018 | 0.68024 | 0.20032 | 1.3852 | -0.0028601 | 0.10716 | 0.0063313 | 0.0020975 | 0.0017368 | 0.087194 | 0.00031364 | 0.0057472 | 1.1165 | 0.24326 | 0.038255 | 0.042365 | 0.037633 | 0.035149
8016534 | 0.01351 | 0.014029 | 11.3384 | 0.31696 | 4.3024 | 1.8757 | 0.00084292 | 0.084109 | 0.044641 | 0.01978 | 0.0023615 | 0.033433 | 0.00047089 | 0.010166 | 1.1551 | 0.22955 | 0.032608 | 0.024507 | 0.030425 | 0.024403
8016534 | 0.013364 | 0.04341 | 11.1583 | 0.35938 | 4.441 | 1.9134 | 0.00044597 | 0.085475 | 0.039968 | 0.021087 | 0.002368 | 0.033159 | 0.000521 | 0.0073699 | 1.1191 | 0.22562 | 0.031523 | 0.025142 | 0.031719 | 0.023979
4022297 | 0.0094079 | 0.085809 | 19.1592 | 0.53132 | 1.4068 | 1.3427 | 0.0012431 | 0.18548 | 0.017151 | 0.0027971 | 0.0083939 | 0.066059 | 0.00013596 | 0.0080474 | 1.2831 | 5.6089 | 0.081692 | 0.080091 | 0.083839 | 0.076538
4022297 | 0.0092364 | 0.092791 | 19.2153 | 0.5306 | 1.4291 | 1.3639 | 0.00080633 | 0.19034 | 0.013184 | 0.0030261 | 0.0081605 | 0.065654 | 0.00010706 | 0.008974 | 1.313 | 5.628 | 0.080053 | 0.083352 | 0.08236 | 0.075907
8011583 | 0.008632 | 0.041482 | 12.1957 | 0.31543 | 2.7562 | 1.2602 | 0.00038177 | 0.28092 | 0.022365 | 0.0027776 | 0.0058876 | 0.027134 | 0.00076921 | 0.0075671 | 1.1347 | 1.2344 | 0.018887 | 0.014887 | 0.018769 | 0.013536
8011583 | 0.0086486 | 0.042616 | 12.0199 | 0.30271 | 2.7482 | 1.2422 | 0.00050596 | 0.2816 | 0.051644 | 0.0024678 | 0.006351 | 0.026116 | 0.0007747 | 0.0079754 | 1.0814 | 1.2352 | 0.019177 | 0.013831 | 0.016875 | 0.013375
3022439 | 0.021571 | 0.08035 | 31.0319 | 1.4161 | 3.1429 | 3.156 | 0.0039646 | 0.15078 | 0.11621 | 0.0044166 | 0.0027498 | 0.092589 | 0.00018735 | 0.021014 | 1.1299 | 0.13782 | 0.043664 | 0.038027 | 0.044104 | 0.035742
3022439 | 0.020915 | 0.079731 | 32.3189 | 1.4418 | 3.1388 | 3.1922 | 0.0036219 | 0.15062 | 0.12196 | 0.0037232 | 0.0029139 | 0.09253 | 0.00016392 | 0.018521 | 1.1165 | 0.14426 | 0.045551 | 0.037759 | 0.042653 | 0.037815
8010730 | 0.0049659 | 0.0068863 | 13.7058 | 0.148 | 2.8571 | 0.29865 | -0.00038376 | 0.20084 | 0.017874 | 0.0015647 | 0.0047897 | 0.021083 | 0.00010633 | 0.0080741 | 0.39091 | 0.82062 | 0.011094 | 0.0074478 | 0.010101 | 0.0076981
8010730 | 0.0052325 | 0.0064393 | 14.5988 | 0.14163 | 2.9067 | 0.29236 | -0.00028684 | 0.20291 | 0.019457 | 0.0014108 | 0.0047339 | 0.021573 | 0.00012503 | 0.0062539 | 0.38926 | 0.81029 | 0.011 | 0.0070267 | 0.0096258 | 0.0073115
4014103 | 0.0013745 | 0.0073002 | 1.5754 | 0.095872 | 0.54904 | 0.37844 | -0.00098212 | 0.012605 | 0.0031646 | 0.0011453 | 0.00076565 | 0.0070097 | 4.3258e-05 | 0.0035601 | 0.17854 | 0.046836 | 0.0017567 | 0.00050241 | 0.0014596 | 0.00070064
4014103 | 0.0016333 | 0.0052497 | 1.5707 | 0.092886 | 0.55176 | 0.38069 | -0.00061222 | 0.013022 | 0.0029395 | 0.0012385 | 0.00071409 | 0.0051461 | 5.0576e-05 | 0.0022165 | 0.18253 | 0.048276 | 0.001102 | 0.00085253 | 0.0014689 | 0.00048466
6015484 | 0.015509 | 0.031309 | 22.9412 | 0.47991 | 3.5992 | 1.3776 | 0.0016643 | 0.11556 | 0.018773 | 0.0036314 | 0.0023758 | 0.031057 | 0.00035477 | 0.0069839 | 2.0841 | 0.19209 | 0.042735 | 0.039612 | 0.042701 | 0.037943
6015484 | 0.014971 | 0.034126 | 22.5162 | 0.47782 | 3.607 | 1.3959 | 0.0015403 | 0.1151 | 0.020621 | 0.0037704 | 0.0023299 | 0.030501 | 0.0003062 | 0.0068726 | 2.101 | 0.19972 | 0.042261 | 0.03977 | 0.04111 | 0.037402
7014783 | 0.011274 | 0.029553 | 20.3989 | 0.3548 | 2.2623 | 2.3251 | 0.00021502 | 0.16096 | 0.01882 | 0.0038619 | 0.0058764 | 0.027552 | 0.00012734 | 0.0088468 | 0.89856 | 0.82458 | 0.029155 | 0.026425 | 0.028873 | 0.025244
7014783 | 0.011303 | 0.028068 | 21.4307 | 0.35548 | 2.3008 | 2.3463 | 0.00054722 | 0.16348 | 0.019001 | 0.003726 | 0.0057235 | 0.028813 | 9.0934e-05 | 0.0061968 | 0.91014 | 0.82968 | 0.028089 | 0.026168 | 0.028018 | 0.025059
6015603 | 0.020011 | 0.024577 | 19.2237 | 0.79102 | 4.2273 | 3.2468 | 0.00062328 | 0.13723 | 0.042704 | 0.0026711 | 0.0036165 | 0.040303 | 0.00025895 | 0.0066777 | 1.7671 | 0.1503 | 0.014122 | 0.0072622 | 0.013373 | 0.0081093
6015603 | 0.019277 | 0.020778 | 18.6581 | 0.7659 | 4.1432 | 3.1971 | 0.0004463 | 0.13356 | 0.044482 | 0.0023163 | 0.0033095 | 0.039053 | 0.00025192 | 0.0068279 | 1.7633 | 0.15582 | 0.012322 | 0.007122 | 0.012294 | 0.0076106
3014800 | 0.0057264 | 0.020648 | 10.1152 | 0.26804 | 0.51334 | 1.5676 | 0.0031769 | 0.12318 | 0.0098087 | 0.002827 | 0.0046165 | 0.013006 | 7.1821e-05 | 0.0087346 | 0.79604 | 0.31505 | 0.011943 | 0.010067 | 0.010491 | 0.0089289
3014800 | 0.0058847 | 0.037986 | 9.9864 | 0.26868 | 0.52822 | 1.558 | 0.0029658 | 0.12359 | 0.0099466 | 0.0029201 | 0.0044857 | 0.012933 | 6.8711e-05 | 0.037045 | 0.78939 | 0.31106 | 0.011139 | 0.0090066 | 0.010932 | 0.008278
7010966 | 0.0073186 | 0.015633 | 8.5023 | 0.61259 | 1.5512 | 1.0995 | 0.0030107 | 0.23851 | 0.024173 | 0.0017007 | 0.0058309 | 0.03022 | 0.00020493 | 0.0072359 | 1.5861 | 0.31027 | 0.049481 | 0.048133 | 0.049896 | 0.04736
7010966 | 0.0079663 | 0.015371 | 8.9564 | 0.63189 | 1.5713 | 1.1147 | 0.0037871 | 0.24333 | 0.11314 | 0.0014941 | 0.0062827 | 0.030646 | 0.00018291 | 0.023513 | 1.5526 | 0.31356 | 0.050088 | 0.049256 | 0.047678 | 0.047544
3012743 | 0.003876 | 0.010447 | 0.74508 | 0.12885 | 0.53974 | 0.24051 | -0.00071053 | 0.043611 | 0.012848 | 0.0004011 | 0.0017543 | 0.0043734 | 4.5387e-05 | 0.0032051 | 0.36036 | 0.1178 | 0.0044778 | 0.0026524 | 0.0032278 | 0.002565
3012743 | 0.0036977 | 0.010227 | 0.73977 | 0.12859 | 0.53657 | 0.23614 | -0.00077899 | 0.043072 | 0.012147 | 0.00017939 | 0.0017611 | 0.0044953 | 4.029e-05 | 0.0032105 | 0.3465 | 0.11539 | 0.0029176 | 0.003122 | 0.0036837 | 0.0030089
8011540 | 0.016599 | 0.0053866 | 12.8741 | 0.32166 | 3.5594 | 0.59632 | 0.0005084 | 0.12314 | 0.041991 | 0.0034519 | 0.0026689 | 0.022761 | 0.00067439 | 0.0052251 | 1.0817 | 0.12969 | 0.010613 | 0.0051516 | 0.010446 | 0.0043849
8011540 | 0.021321 | 0.011125 | 23.9723 | 0.33526 | 3.5922 | 0.58931 | 0.0018639 | 0.11834 | 0.042974 | 0.0028134 | 0.0026239 | 0.023709 | 0.00063935 | 0.009061 | 1.106 | 0.125 | 0.010544 | 0.0047639 | 0.011417 | 0.0050489
8019240 | 0.013519 | 0.042221 | 13.227 | 0.38743 | 5.0159 | 2.3699 | 0.0029121 | 0.40038 | 0.02747 | 0.0047782 | 0.0062857 | 0.030699 | 0.00051322 | 0.091436 | 1.6094 | 0.37876 | 0.033577 | 0.028617 | 0.034866 | 0.028661
8019240 | 0.01166 | 0.034499 | 5.994 | 0.37597 | 4.96 | 2.3402 | -9.5038e-05 | 0.39989 | 0.026661 | 0.0043775 | 0.0062876 | 0.029091 | 0.00047965 | 0.088684 | 1.6167 | 0.37333 | 0.033154 | 0.028514 | 0.034639 | 0.028237
4013190 | 0.0066947 | 0.0040842 | 5.2796 | 0.15424 | 1.1684 | 0.78055 | -0.0005061 | 0.063809 | 0.0085052 | 0.00041877 | 0.0038323 | 0.015735 | 6.6678e-05 | 0.0027391 | 0.74742 | 0.24138 | 0.01104 | 0.0077495 | 0.010678 | 0.0081812
4013190 | 0.0066362 | 0.0042771 | 5.2576 | 0.15416 | 1.1573 | 0.78455 | -0.00039982 | 0.064387 | 0.016952 | 0.00054175 | 0.0036969 | 0.016153 | 6.7877e-05 | 0.0030026 | 0.78442 | 0.25276 | 0.012482 | 0.0077011 | 0.010402 | 0.0083082
5012333 | 0.0091659 | 0.039559 | 10.8423 | 0.18847 | 2.7323 | 0.33572 | 0.0028312 | 0.099121 | 0.024399 | 0.00058211 | 0.0025841 | 0.019008 | 0.00010711 | 0.0064146 | 0.58543 | 0.24911 | 0.016424 | 0.01358 | 0.014866 | 0.012857
8018855 | 0.007528 | 0.036497 | 32.4768 | 0.2852 | 2.2395 | 1.7761 | 0.0019473 | 0.18974 | 0.012093 | 0.013892 | 0.0030116 | 0.010306 | 0.00067782 | 0.0032869 | 1.2947 | 0.51586 | 0.013715 | 0.0092619 | 0.011701 | 0.0084225
4016084 | 0.0076211 | 0.023935 | 7.6111 | 0.33845 | 0.95379 | 0.80039 | 0.0013678 | 0.24183 | 0.012283 | 0.0010051 | 0.011181 | 0.02117 | 0.00019561 | 0.0033201 | 1.7818 | 0.14523 | 0.017298 | 0.013999 | 0.01539 | 0.013107
5011183 | 0.015057 | 0.061802 | 8.4708 | 0.5193 | 0.926 | 1.0341 | 0.0016604 | 0.16217 | 0.01271 | 0.00085226 | 0.0034534 | 0.033914 | 0.00012463 | 0.0050676 | 1.2258 | 0.12892 | 0.015763 | 0.012722 | 0.014023 | 0.011638
5021782 | 0.0063412 | 0.027369 | 5.952 | 0.2601 | 2.781 | 0.68474 | 0.00095642 | 0.076807 | 0.053981 | 0.00090306 | 0.0030576 | 0.014433 | 5.9739e-05 | 0.0012281 | 0.83211 | 1.8864 | 0.012565 | 0.0097171 | 0.011715 | 0.009026
5019109 | 0.0078114 | 0.13048 | 15.584 | 0.44922 | 2.9408 | 2.5472 | 0.0042576 | 0.11063 | 0.061661 | 0.002211 | 0.0031098 | 0.017481 | 0.00014525 | 0.011311 | 1.6501 | 0.36496 | 0.012324 | 0.0086394 | 0.011075 | 0.0063366
6015450 | 0.0035789 | 0.010711 | 3.266 | 0.090394 | 0.84943 | 0.50049 | 0.00059378 | 0.036354 | 0.0066739 | 0.00048701 | 0.0014906 | 0.01223 | 2.0953e-05 | -0.00040582 | 0.22432 | 0.077194 | 0.0059128 | 0.0041074 | 0.0043354 | 0.0039779
3019454 | 0.0032295 | 0.033773 | 15.4588 | 0.15542 | 0.24223 | 0.27711 | 0.00067703 | 0.037655 | 0.003767 | 0.00071964 | 0.0010372 | 0.0060085 | 7.435e-05 | 0.0014336 | 0.28603 | 0.11751 | 0.014127 | 0.014211 | 0.013815 | 0.012792
3010465 | 0.034314 | 0.10926 | 12.8417 | 0.48075 | 2.4391 | 0.73852 | 0.0025815 | 0.21017 | 0.072513 | 0.00039549 | 0.0033898 | 0.0098877 | 4.6881e-05 | 0.026284 | 1.1081 | 0.33838 | 0.0064471 | 0.0036066 | 0.0057965 | 0.0035765
5010926 | 0.0078054 | 0.037938 | 19.0068 | 0.35091 | 2.6288 | 0.36277 | 0.00086207 | 0.083265 | 0.028187 | 0.0004884 | 0.0034356 | 0.0050027 | 9.0392e-05 | 0.0040822 | 4.2059 | 3.6942 | 0.012016 | 0.013764 | 0.012402 | 0.010006
7010788 | 0.0053346 | 0.017756 | 3.3825 | 0.14237 | 0.76073 | 2.1262 | 0.00035921 | 0.045869 | 0.007681 | 0.00033096 | 0.0029693 | 0.0093822 | 5.8025e-05 | 0.0022379 | 0.49071 | 0.19589 | 0.0078257 | 0.0054875 | 0.0072923 | 0.0056782
7010788 | 0.0059942 | 0.017849 | 3.2474 | 0.1418 | 0.74983 | 2.1081 | 0.00031485 | 0.045507 | 0.0106 | 0.00040248 | 0.0027114 | 0.0095544 | 3.0653e-05 | 0.0017494 | 0.49481 | 0.19319 | 0.0076343 | 0.0054628 | 0.0064683 | 0.0053832
7042000 | 0.0099525 | 0.095753 | 7.1623 | 1.5219 | 2.9522 | 1.6813 | 0.0022922 | 0.15195 | 0.12905 | 0.0025944 | 0.024918 | 0.020818 | 0.00015431 | 0.0037948 | 0.79039 | 0.49309 | 0.01744 | 0.012569 | 0.0159 | 0.012371
7042000 | 0.010568 | 0.099281 | 7.1228 | 1.5352 | 2.8878 | 1.6469 | 0.0026175 | 0.15111 | 0.12879 | 0.0027043 | 0.024663 | 0.02372 | 0.00015593 | 0.004567 | 0.78678 | 0.48605 | 0.017062 | 0.012569 | 0.016021 | 0.01269
8017425 | 0.011348 | 0.1844 | 5.8445 | 0.38097 | 3.4798 | 0.77277 | 0.00046975 | 0.52912 | 0.06393 | 0.00072449 | 0.00571 | 0.015327 | 0.00019395 | 0.0017028 | 1.2648 | 0.78376 | 0.014207 | 0.0091576 | 0.013043 | 0.0091332
8017425 | 0.011932 | 0.19103 | 5.8094 | 0.39027 | 3.4941 | 0.76507 | 0.00057173 | 0.52713 | 0.054909 | 0.00081004 | 0.005189 | 0.015719 | 0.00021679 | 0.0020053 | 1.2576 | 0.78395 | 0.013121 | 0.0095329 | 0.01258 | 0.0089458
8022810 | 0.0020717 | 0.014988 | 1.4971 | 0.091641 | 1.2924 | 0.23329 | 0.0010313 | 0.051716 | 0.019637 | 0.0037722 | 0.0023714 | 0.0071425 | 8.402e-05 | 0.0042224 | 0.27679 | 0.77671 | 0.005457 | 0.0040393 | 0.004712 | 0.0042305
8022810 | 0.0020916 | 0.018256 | 1.5783 | 0.093848 | 1.2967 | 0.23054 | 0.00054174 | 0.051733 | 0.022584 | 0.0034976 | 0.0022847 | 0.0089246 | 0.00011925 | 0.0046114 | 0.27366 | 0.78718 | 0.0054394 | 0.0044562 | 0.0047175 | 0.0041827
6022170 | 0.010159 | 0.089608 | 16.6361 | 0.40361 | 3.25 | 1.3449 | 0.0015387 | 0.22078 | 0.066967 | 0.0018035 | 0.0019004 | 0.018192 | 7.2227e-05 | 0.030573 | 1.8837 | 0.1156 | 0.0078945 | 0.0077123 | 0.0067098 | 0.003632
6022170 | 0.010186 | 0.088253 | 16.8718 | 0.37424 | 3.3146 | 1.3887 | 0.0011206 | 0.2262 | 0.066991 | 0.0022187 | 0.0017701 | 0.013955 | 8.6997e-05 | 0.029323 | 1.9504 | 0.12085 | 0.0075763 | 0.003369 | 0.0063266 | 0.0033226
6018866 | 0.0022914 | 0.014206 | 15.9737 | 0.11833 | 0.23443 | 0.83726 | 0.00016061 | 0.04173 | 0.0079791 | 0.00087253 | 0.0010208 | 0.0060964 | 5.7157e-05 | 0.0090733 | 0.32343 | 0.028837 | 0.0032821 | 0.0014005 | 0.0025572 | 0.0015738
6018866 | 0.0022535 | 0.016443 | 28.9826 | 0.1248 | 0.24597 | 0.8654 | 0.00040943 | 0.042526 | 0.011053 | 0.0010771 | 0.0011245 | 0.007063 | 7.6793e-05 | 0.012709 | 0.32042 | 0.027499 | 0.0032372 | 0.0020307 | 0.002711 | 0.00171
6015166 | 0.007784 | 0.13904 | 7.8799 | 0.28804 | 1.8187 | 0.67364 | 0.00077322 | 0.055596 | 0.016985 | 0.00075082 | 0.0012538 | 0.0086131 | 0.00020862 | 0.0065974 | 0.57656 | 0.064778 | 0.015075 | 0.012439 | 0.014393 | 0.012353
6015166 | 0.0079784 | 0.14381 | 8.0733 | 0.29704 | 1.8484 | 0.68536 | 0.00088232 | 0.05548 | 0.019024 | 0.0010856 | 0.0011955 | 0.010927 | 0.00022527 | 0.0087352 | 0.5896 | 0.066681 | 0.015982 | 0.012697 | 0.013834 | 0.01175
3010007 | 0.006087 | 0.05135 | 14.0791 | 0.30145 | 3.4617 | 1.2625 | 0.0010166 | 0.091456 | 0.01626 | 0.002132 | 0.0056695 | 0.010299 | 9.5525e-05 | 0.0076438 | 1.0775 | 1.0576 | 0.012201 | 0.0083851 | 0.010297 | 0.0076514
3010007 | 0.00598 | 0.05341 | 14.4507 | 0.31061 | 3.5501 | 1.2984 | 0.0010603 | 0.092747 | 0.016758 | 0.0021193 | 0.0057541 | 0.01063 | 0.00013789 | 0.010259 | 1.125 | 1.0945 | 0.012884 | 0.007884 | 0.010548 | 0.0084087
3015912 | 0.0051512 | 0.029558 | 11.2056 | 0.24188 | 1.0186 | 0.60291 | 0.00055536 | 0.071505 | 0.022355 | 0.00086381 | 0.0018508 | 0.011796 | 4.379e-05 | 0.0064534 | 0.48486 | 0.052632 | 0.0097005 | 0.0067517 | 0.0086268 | 0.006248
3015912 | 0.0047765 | 0.02923 | 11.1645 | 0.23482 | 1.0143 | 0.59634 | 0.00045982 | 0.069651 | 0.015831 | 0.00072423 | 0.0019017 | 0.012107 | 4.4322e-05 | 0.0050215 | 0.46136 | 0.053733 | 0.0098755 | 0.0073422 | 0.0086997 | 0.0066719
4021312 | 0.0013338 | 0.0075429 | 3.1535 | 0.067894 | 0.21769 | 0.30654 | -9.5029e-05 | 0.034207 | 0.0036692 | 0.00022555 | 0.00082547 | 0.003385 | 3.4228e-05 | 0.00046111 | 0.30734 | 0.071493 | 0.0020137 | 0.00094063 | 0.0016336 | 0.00086337
4021312 | 0.0015068 | 0.008627 | 3.0255 | 0.066062 | 0.21039 | 0.289 | -9.5411e-05 | 0.032519 | 0.0037083 | 0.00024276 | 0.00074178 | 0.0033553 | 1.6735e-05 | 0.00077163 | 0.29321 | 0.066255 | 0.00181 | 0.00078303 | 0.001571 | 0.00094027
5017041 | 0.0049288 | 0.018091 | 16.0834 | 0.30108 | 1.923 | 0.22683 | 0.0089634 | 0.084776 | 0.020423 | 0.0014175 | 0.0026861 | 0.027627 | 0.0019449 | 0.0015324 | 0.94497 | 0.314 | 0.0078191 | 0.0062119 | 0.0078956 | 0.0055271
5017041 | 0.0054905 | 0.018318 | 15.6981 | 0.29813 | 1.8611 | 0.21845 | 0.0096075 | 0.081002 | 0.016765 | 0.00094783 | 0.0025157 | 0.026344 | 0.0017651 | 0.0011028 | 0.91724 | 0.31075 | 0.0081867 | 0.0058203 | 0.0069498 | 0.0056396
5016142 | 0.0028281 | 0.013741 | 5.3706 | 0.13919 | 0.49685 | 0.42325 | 0.00010396 | 0.046704 | 0.013842 | 8.817e-05 | 0.0032213 | 0.0048583 | 1.0635e-05 | 0.00096611 | 0.70724 | 0.15658 | 0.0056317 | 0.0052757 | 0.0057995 | 0.0049728
5016142 | 0.0029252 | 0.014794 | 5.5074 | 0.14741 | 0.50343 | 0.42983 | -0.00017895 | 0.047286 | 0.013071 | 0.00010659 | 0.0028114 | 0.0054304 | 3.4905e-06 | 0.0011397 | 0.74629 | 0.15673 | 0.0065075 | 0.004885 | 0.0054796 | 0.0049552
5016940 | 0.0090914 | 0.077832 | 15.9401 | 0.30144 | 2.0315 | 1.1945 | 0.001391 | 0.10728 | 0.066794 | 0.00093519 | 0.002772 | 0.025441 | 7.3707e-05 | 0.0097762 | 1.0182 | 0.11537 | 0.024786 | 0.020857 | 0.022527 | 0.019913
5016940 | 0.0094001 | 0.084101 | 16.141 | 0.31104 | 2.0568 | 1.2306 | 0.0014657 | 0.10813 | 0.13568 | 0.00082308 | 0.0031317 | 0.02522 | 9.0165e-05 | 0.010467 | 1.043 | 0.11865 | 0.023489 | 0.021065 | 0.023269 | 0.021325
5018021 | 0.014024 | 0.051838 | 7.4613 | 0.29536 | 4.1146 | 0.63656 | 4.1786e-05 | 0.20179 | 0.82991 | 0.00064134 | 0.0016973 | 0.015484 | 3.4252e-05 | 0.0039858 | 1.1943 | 0.29274 | 0.0076138 | 0.0032926 | 0.0055075 | 0.0036247
5018021 | 0.013427 | 0.054494 | 9.7604 | 0.29328 | 4.1281 | 0.64165 | 0.00038875 | 0.20082 | 0.8213 | 0.00051541 | 0.0017527 | 0.015771 | 5.6036e-05 | 0.0046205 | 1.1679 | 0.29176 | 0.0075441 | 0.0029127 | 0.0063056 | 0.0032421
6014410 | 0.0052054 | 0.027796 | 9.6639 | 0.1421 | 1.5693 | 0.57807 | 0.00033694 | 0.061823 | 0.048019 | 0.00048846 | 0.0010515 | 0.0049801 | 2.5108e-05 | 0.0020409 | 0.63811 | 0.049005 | 0.0049941 | 0.0022463 | 0.00462 | 0.0021609
6014410 | 0.0052983 | 0.026149 | 9.6097 | 0.14357 | 1.57 | 0.57952 | 0.00083827 | 0.060764 | 0.046912 | 0.0004401 | 0.001135 | 0.0050321 | 4.4284e-05 | 0.012469 | 0.62621 | 0.050003 | 0.0052754 | 0.0025455 | 0.0041522 | 0.0023256
6010377 | 0.016625 | 0.12547 | 39.8477 | 0.59653 | 1.3876 | 7.8132 | 0.0052062 | 0.13079 | 0.019004 | 0.0041824 | 0.0066939 | 0.088989 | 0.00054091 | 0.50571 | 5.8169 | 0.11811 | 0.086295 | 0.085122 | 0.083004 | 0.084265
6010377 | 0.016485 | 0.12996 | 40.3364 | 0.59988 | 1.3994 | 8.0247 | 0.005532 | 0.12834 | 0.018736 | 0.0045893 | 0.0061957 | 0.088052 | 0.00041708 | 0.26711 | 5.8472 | 0.12271 | 0.084826 | 0.083535 | 0.082557 | 0.083488
8015473 | 0.010624 | 0.074335 | 27.0768 | 0.53658 | 2.3097 | 2.1271 | 0.0023653 | 0.16829 | 0.037881 | 0.0022291 | 0.0041009 | 0.023285 | 0.002154 | 0.0027742 | 2.3729 | 0.30562 | 0.042024 | 0.040077 | 0.041147 | 0.038862
8015473 | 0.010127 | 0.07655 | 27.4878 | 0.55104 | 2.3754 | 2.1482 | 0.0022487 | 0.17073 | 0.03772 | 0.0026495 | 0.0042702 | 0.023088 | 0.0021511 | 0.003754 | 2.4329 | 0.30627 | 0.043491 | 0.040582 | 0.041151 | 0.038689
3011437 | 0.026478 | 0.035569 | 20.224 | 0.29839 | 1.1034 | 0.40797 | 0.0016768 | 0.096816 | 0.031122 | 0.00025702 | 0.0035965 | 0.012939 | 5.9361e-05 | 0.0017244 | 2.7424 | 0.2062 | 0.012655 | 0.0092217 | 0.011562 | 0.0091685
3012395 | 0.0052132 | 0.016524 | 3.5651 | 0.10995 | 1.2537 | 0.39861 | 9.2173e-05 | 0.049303 | 0.040519 | 0.00047127 | 0.00803 | 0.0062372 | 0.00014512 | 0.0030811 | 0.40813 | 0.053722 | 0.010289 | 0.00861 | 0.010188 | 0.0079974
3012395 | 0.0042472 | 0.015632 | 3.5874 | 0.10196 | 1.195 | 0.37047 | 1.7011e-05 | 0.047181 | 0.040415 | 0.00024872 | 0.0075838 | 0.005896 | 0.00015372 | 0.0050593 | 0.38545 | 0.049894 | 0.011785 | 0.0093267 | 0.010669 | 0.008041
4015282 | 0.006369 | 0.023513 | 5.6169 | 0.19956 | 1.4023 | 0.52973 | 0.0037055 | 0.10137 | 0.02758 | 0.0011164 | 0.0022209 | 0.040555 | 0.00020647 | 0.0050528 | 0.5838 | 0.093868 | 0.0079868 | 0.0047003 | 0.0068809 | 0.00366
4015282 | 0.0060334 | 0.022337 | 5.6356 | 0.20031 | 1.4268 | 0.53788 | 0.0031816 | 0.098047 | 0.028368 | 0.0010826 | 0.0022595 | 0.041062 | 0.00017684 | 0.0030659 | 0.59139 | 0.092507 | 0.0079177 | 0.0036492 | 0.0064026 | 0.0032688
4011724 | 0.0052584 | 0.024371 | 7.729 | 0.20471 | 1.2453 | 0.39543 | 0.0006493 | 0.068671 | 0.0099062 | 0.00069886 | 0.0028019 | 0.0068653 | 8.079e-05 | -0.00035525 | 0.32836 | 0.28909 | 0.014382 | 0.011421 | 0.014469 | 0.010833
4011724 | 0.0061554 | 0.023679 | 7.7925 | 0.20203 | 1.255 | 0.36823 | 0.00010121 | 0.074402 | 0.0072029 | 0.00039927 | 0.0030991 | 0.0060325 | 6.2901e-05 | -0.00060186 | 0.33997 | 0.30403 | 0.01368 | 0.011673 | 0.014206 | 0.011227
4013271 | 0.0068475 | 0.015281 | 13.1147 | 0.13216 | 1.1575 | 1.7863 | 0.00023221 | 0.051084 | 0.060016 | 0.0014205 | 0.0025085 | 0.0040817 | 7.3343e-05 | -0.0012629 | 1.496 | 0.79352 | 0.0068631 | 0.0046457 | 0.0060973 | 0.0040873
4013271 | 0.0068458 | 0.015626 | 13.1884 | 0.13503 | 1.1696 | 1.8129 | 0.00045364 | 0.05022 | 0.060336 | 0.0010652 | 0.0024317 | 0.003936 | 5.4063e-05 | -0.0015258 | 1.4922 | 0.79382 | 0.0064545 | 0.0048414 | 0.006489 | 0.0044633
6022359 | 0.0057768 | 0.028262 | 13.9472 | 0.21261 | 1.6336 | 0.7033 | 0.00042903 | 0.076041 | 0.015332 | 0.00065717 | 0.0026966 | 0.030205 | 8.082e-05 | -0.00022716 | 0.97087 | 0.090143 | 0.0075799 | 0.0051957 | 0.0069016 | 0.0044588
6022359 | 0.0056078 | 0.030524 | 14.2063 | 0.27503 | 1.6655 | 0.70069 | 0.00073274 | 0.076997 | 0.018391 | 0.0008016 | 0.002622 | 0.030442 | 3.8037e-05 | 0.0021438 | 0.97797 | 0.093335 | 0.0070384 | 0.004981 | 0.0062155 | 0.0045955
7014953 | 0.010544 | 0.067293 | 5.3134 | 0.25866 | 4.12 | 0.96692 | 0.00041549 | 0.2003 | 0.060199 | 0.00075813 | 0.0072302 | 0.034167 | 4.767e-05 | 0.0014113 | 0.69831 | 0.34029 | 0.012449 | 0.0070464 | 0.012464 | 0.0070637
7014953 | 0.01017 | 0.062809 | 5.2467 | 0.26464 | 4.0366 | 0.96911 | 0.0004054 | 0.18902 | 0.056534 | 0.000668 | 0.0067861 | 0.034792 | 3.5912e-05 | 0.00055567 | 0.68894 | 0.33314 | 0.013816 | 0.0065574 | 0.012861 | 0.0072524
7016182 | 0.0096316 | 0.022832 | 12.7279 | 0.26207 | 1.9192 | 1.0762 | 0.00095411 | 0.12752 | 0.025794 | 0.00069021 | 0.0074123 | 0.017375 | 8.0931e-05 | 0.001 | 0.86019 | 0.26753 | 0.018323 | 0.016985 | 0.017882 | 0.015373
7016182 | 0.0095013 | 0.021264 | 12.2099 | 0.27141 | 1.8722 | 1.0692 | 0.00067382 | 0.12165 | 0.022574 | 0.00084031 | 0.0070617 | 0.015523 | 7.3387e-05 | 0 | 0.86907 | 0.26821 | 0.018221 | 0.015669 | 0.017027 | 0.015222
8016577 | 0.018112 | 0.04925 | 28.3113 | 0.50763 | 7.5805 | 1.0441 | 0.0038975 | 0.12567 | 0.081456 | 0.0015258 | 0.0016542 | 0.025572 | 0.00076535 | 0.053 | 1.6507 | 0.1571 | 0.016992 | 0.011486 | 0.017016 | 0.011588
8016577 | 0.018291 | 0.052073 | 24.7478 | 0.5061 | 7.4458 | 1.0281 | 0.0039611 | 0.12338 | 0.081864 | 0.0012972 | 0.0016191 | 0.025186 | 0.00052509 | 0.044 | 1.6299 | 0.15428 | 0.016642 | 0.011283 | 0.016942 | 0.010798
8015058 | 0.015514 | 0.11139 | 28.0344 | 0.66299 | 5.5335 | 1.1639 | 0.0016496 | 0.58389 | 0.047304 | 0.0016647 | 0.028592 | 0.04407 | 0.0004789 | 0.014 | 2.7634 | 0.90773 | 0.026276 | 0.022221 | 0.026434 | 0.021188
8015058 | 0.015713 | 0.10504 | 27.6372 | 0.65965 | 5.6895 | 1.1762 | 0.0015268 | 0.56444 | 0.048896 | 0.0022032 | 0.028291 | 0.043348 | 0.00040403 | 0.011 | 2.7668 | 0.91108 | 0.025396 | 0.022019 | 0.026329 | 0.021454
8010064 | 0.0078706 | 0.029737 | 6.6738 | 0.4269 | 3.3765 | 1.4074 | 0.0012347 | 0.20425 | 0.020169 | 0.0017173 | 0.0077062 | 0.014267 | 0.00071039 | 0 | 1.4675 | 1.1028 | 0.048048 | 0.04459 | 0.049573 | 0.044353
8010064 | 0.0080906 | 0.028377 | 5.9425 | 0.43481 | 3.3148 | 1.3912 | 0.00096946 | 0.2 | 0.019147 | 0.0016147 | 0.0076031 | 0.01377 | 0.00069105 | -0.001 | 1.5051 | 1.1019 | 0.047746 | 0.044304 | 0.049748 | 0.046103
8010226 | 0.0050348 | 0.023888 | 16.0696 | 0.22947 | 2.5077 | 0.84543 | 0.001625 | 0.10189 | 0.023152 | 0.00479 | 0.0028473 | 0.0081677 | 0.00090694 | -0.001 | 0.90463 | 0.42737 | 0.010782 | 0.0079555 | 0.010831 | 0.0085342
8010226 | 0.0050929 | 0.020228 | 16.1402 | 0.23028 | 2.4929 | 0.8522 | 0.0013489 | 0.10198 | 0.02068 | 0.005344 | 0.0028474 | 0.0069329 | 0.00092429 | -0.001 | 0.89897 | 0.44151 | 0.0098063 | 0.0076517 | 0.0099681 | 0.0082228
8016984 | 0.0068462 | 0.02647 | 5.8112 | 0.33725 | 2.1483 | 0.63949 | 0.0008833 | 0.16122 | 0.013847 | 0.0060076 | 0.0041458 | 0.027095 | 0.001492 | -0.001 | 1.2426 | 0.18711 | 0.013168 | 0.0089018 | 0.012243 | 0.0078074
8016984 | 0.0063977 | 0.028826 | 6.1212 | 0.41186 | 2.1858 | 0.64892 | 0.0012152 | 0.16171 | 0.013678 | 0.0063053 | 0.0041415 | 0.02778 | 0.0016223 | -0.001 | 1.2495 | 0.18471 | 0.012613 | 0.0096457 | 0.012968 | 0.0074335
3011828 | 0.0057516 | 0.055285 | 10.6643 | 0.26174 | 0.84984 | 1.272 | 0.0011467 | 0.051762 | 0.015807 | 0.0021212 | 0.0037199 | 0.0040449 | 7.75e-05 | 0.17075 | 0.59704 | 1.4285 | 0.018 | 0.018 | 0.02 | 0.016
3011828 | 0.0049663 | 0.035225 | 9.7928 | 0.22369 | 0.78488 | 1.2021 | 0.0014299 | 0.048832 | 0.013461 | 0.0016095 | 0.0033964 | 0.0034212 | 8.4551e-05 | 0.15493 | 0.54748 | 1.2998 | 0.017 | 0.017 | 0.017 | 0.016
3010848 | 0.0028106 | 0.0090126 | 5.7358 | 0.078347 | 0.75758 | 0.42961 | 0.00065635 | 0.026622 | 0.006426 | 0.00028204 | 0.0014515 | 0.005484 | 0.00015695 | 0.005082 | 0.35882 | 0.026054 | 0.005 | 0.004 | 0.005 | 0.004
3010848 | 0.0025989 | 0.0093539 | 5.3468 | 0.073163 | 0.71866 | 0.40444 | 0.00055232 | 0.025672 | 0.0050205 | 0.00031273 | 0.0015594 | 0.0046685 | 0.00016873 | 0.0016206 | 0.34523 | 0.025061 | 0.004 | 0.004 | 0.004 | 0.004
3018350 | 0.0031423 | 0.012486 | 1.2836 | 0.077382 | 0.83059 | 0.22698 | 0.00090091 | 0.040096 | 0.021827 | 0.00028507 | 0.0062177 | 0.0042482 | 1.8382e-05 | 0.0087435 | 0.20688 | 0.094307 | 0.003 | 0.004 | 0.004 | 0.003
3018350 | 0.0028537 | 0.012147 | 1.323 | 0.081566 | 0.83435 | 0.23416 | 0.0013236 | 0.04031 | 0.023029 | 0.0002416 | 0.0063393 | 0.0062766 | 1.5022e-05 | 0.020166 | 0.20276 | 0.091634 | 0.004 | 0.004 | 0.004 | 0.003
3013995 | 0.004608 | 0.020298 | 12.2948 | 0.32517 | 1.1951 | 0.46117 | 0.00074664 | 0.13542 | 0.016272 | 0.00075997 | 0.0056447 | 0.016308 | 5.1274e-05 | 0.0027234 | 0.9995 | 0.091373 | 0.029 | 0.025 | 0.028 | 0.026
3013995 | 0.005044 | 0.01891 | 12.4932 | 0.33473 | 1.2099 | 0.47256 | 0.0008584 | 0.13743 | 0.0207 | 0.00087966 | 0.0057586 | 0.016888 | 3.5345e-05 | 0.0026418 | 0.98536 | 0.093143 | 0.03 | 0.027 | 0.029 | 0.027
3022056 | 0.0082074 | 0.027224 | 9.69 | 0.081832 | 0.86524 | 0.26156 | 0.00016769 | 0.03286 | 0.012884 | 0.00013256 | 0.0018476 | 0.0014211 | 5.4025e-05 | 0.00287 | 0.26118 | 0.022029 | 0.005 | 0.005 | 0.005 | 0.006
3022056 | 0.0084571 | 0.025908 | 9.5211 | 0.083109 | 0.83803 | 0.25544 | 0.00026534 | 0.032541 | 0.013476 | 0.0004677 | 0.0019526 | 0.00086138 | 4.9822e-05 | 0.0021099 | 0.2676 | 0.021713 | 0.005 | 0.005 | 0.006 | 0.005
4018125 | 0.011897 | 0.077096 | 19.4727 | 1.4964 | 0.69722 | 1.0886 | 0.0012482 | 0.12453 | 0.00488 | 0.00036127 | 0.005886 | 0.08661 | 0.00012718 | -0.0012169 | 1.7539 | 0.22814 | 0.009 | 0.006 | 0.008 | 0.005
4018125 | 0.011468 | 0.084439 | 19.3583 | 1.4994 | 0.70085 | 1.0932 | 0.0014199 | 0.12178 | 0.0041891 | 0.00046115 | 0.0061631 | 0.084882 | 0.00010768 | -0.0015971 | 1.7612 | 0.22652 | 0.009 | 0.006 | 0.009 | 0.005
4015800 | 0.01237 | 0.031002 | 16.1409 | 0.37868 | 3.2468 | 0.66561 | 0.00094269 | 0.17801 | 0.036738 | 0.00086171 | 0.0045029 | 0.011871 | 9.9012e-05 | 0.0027111 | 1.4877 | 0.34427 | 0.037565 | 0.032991 | 0.035767 | 0.032376
4015800 | 0.012948 | 0.029017 | 16.3264 | 0.39261 | 3.2785 | 0.68022 | 0.00092191 | 0.1834 | 0.039403 | 0.0007016 | 0.0048635 | 0.013311 | 0.00013552 | 0.0026651 | 1.5145 | 0.35971 | 0.037467 | 0.031584 | 0.035219 | 0.031486
5013429 | 0.010941 | 0.045687 | 2.0934 | 0.093967 | 0.64874 | 0.3144 | 0.00035858 | 0.095225 | 0.010494 | 0.00083408 | 0.0022614 | 0.010576 | 3.154e-05 | -0.0026292 | 0.36113 | 0.048049 | 0.010086 | 0.0085878 | 0.010181 | 0.0078924
5013429 | 0.011872 | 0.048153 | 1.1533 | 0.09345 | 0.63716 | 0.31919 | 0.00049204 | 0.096396 | 0.023542 | 0.00069502 | 0.0022304 | 0.010336 | 5.9889e-05 | -0.0024626 | 0.3598 | 0.047679 | 0.011492 | 0.008576 | 0.01047 | 0.0078788
5011264 | 0.0081502 | 0.033483 | 16.6419 | 0.2632 | 1.3409 | 0.78761 | 0.0028611 | 0.11328 | 0.041609 | 0.00068881 | 0.0029354 | 0.016405 | 4.4228e-05 | 0.0020929 | 1.4232 | 0.65797 | 0.02567 | 0.024508 | 0.025255 | 0.025238
5011264 | 0.00847 | 0.035305 | 17.0077 | 0.27672 | 1.3768 | 0.80927 | 0.0027926 | 0.11532 | 0.018365 | 0.0005039 | 0.0027206 | 0.017892 | 4.0831e-05 | 0.0016612 | 1.48 | 0.67048 | 0.028678 | 0.025602 | 0.02727 | 0.02458
5014085 | 0.0092607 | 0.24598 | 26.9315 | 0.51171 | 0.53157 | 1.3231 | 0.0015752 | 0.47956 | 0.0031695 | 0.0022091 | 0.0090023 | 0.033722 | 0.00018712 | 0.084902 | 0.92775 | 0.22596 | 0.016885 | 0.013382 | 0.016382 | 0.013432
5014085 | 0.0094019 | 0.24114 | 27.2895 | 0.51784 | 0.53044 | 1.3221 | 0.0013627 | 0.47938 | 0.0038014 | 0.0022872 | 0.0096338 | 0.033826 | 0.00012954 | 0.084057 | 0.95163 | 0.22645 | 0.016081 | 0.014518 | 0.016746 | 0.013351
5010322 | 0.0033977 | 0.016556 | 4.3652 | 0.24217 | 1.0135 | 0.55929 | 0.00064957 | 0.052397 | 0.0077882 | 0.001372 | 0.0036353 | 0.0041954 | 6.0643e-05 | -0.0021332 | 0.81312 | 1.5813 | 0.011755 | 0.0097629 | 0.011033 | 0.0099577
5010322 | 0.0029731 | 0.017238 | 4.3129 | 0.2405 | 1.0005 | 0.55614 | 0.00078788 | 0.051685 | 0.0063307 | 0.0017008 | 0.0037072 | 0.0064807 | 7.5321e-05 | -0.0020192 | 0.80253 | 1.5769 | 0.012158 | 0.010531 | 0.011152 | 0.0099873
5017319 | 0.052562 | 0.028571 | 5.8109 | 0.20255 | 1.4704 | 0.30383 | 0.00086972 | 0.14706 | 0.042779 | 0.00045799 | 0.0080182 | 0.013617 | 4.5153e-05 | -0.0010797 | 0.79138 | 0.24248 | 0.0102 | 0.0064041 | 0.0087863 | 0.0061923
5017319 | 0.052674 | 0.027078 | 5.7655 | 0.2036 | 1.4511 | 0.29834 | 0.00079223 | 0.14611 | 0.043576 | 0.00055713 | 0.0079381 | 0.014023 | 7.8516e-05 | -0.0014003 | 0.80212 | 0.25074 | 0.010028 | 0.0062765 | 0.0094629 | 0.0060731
3014908 | 0.0031503 | 0.029416 | 2.2999 | 0.15649 | 0.68041 | 0.2404 | 0.0010989 | 0.026621 | 0.026335 | 0.00045022 | 0.00081793 | 0.011781 | 0.00023075 | 0.014435 | 0.13271 | 0.028593 | 0.0025917 | 0.0022663 | 0.0018323 | 0.0014278
3014908 | 0.0027949 | 0.051998 | 1.8256 | 0.093738 | 0.66667 | 0.23597 | 0.002905 | 0.024985 | 0.020116 | 0.00010112 | 0.00068528 | 0.0025602 | 0.00011241 | -0.0055065 | 0.13868 | 0.029617 | 0.0014696 | 0.0015602 | 0.0022925 | 0.0016998
6013546 | 0.0044875 | 0.039033 | 5.0972 | 0.50415 | 0.76991 | 0.59169 | 0.0011229 | 0.18585 | 0.015473 | 0.0011379 | 0.003421 | 0.029716 | 8.029e-05 | 0.0040477 | 1.4638 | 0.15605 | 0.01052 | 0.0094289 | 0.010574 | 0.0091873
6013546 | 0.0053074 | 0.041974 | 5.062 | 0.49513 | 0.76658 | 0.59367 | 0.00092554 | 0.18974 | 0.010811 | 0.0010165 | 0.0035156 | 0.028653 | 6.1386e-05 | 0.0041545 | 1.4637 | 0.15761 | 0.01048 | 0.0083972 | 0.0084788 | 0.0082266
7010320 | 0.0064038 | 0.044089 | 10.4066 | 0.24176 | 2.9178 | 0.80288 | 0.0019095 | 0.1581 | 0.30216 | 0.0022678 | 0.0046084 | 0.037849 | 0.00019781 | -0.002811 | 1.1679 | 0.48444 | 0.015644 | 0.011865 | 0.015819 | 0.0099975
7010320 | 0.0072777 | 0.044369 | 12.2588 | 0.22833 | 2.9116 | 0.80893 | 0.0012856 | 0.16167 | 0.31015 | 0.0025616 | 0.0048435 | 0.038485 | 0.00028665 | -0.002515 | 1.1828 | 0.4904 | 0.016188 | 0.011716 | 0.015711 | 0.0091678
7014678 | 0.0054057 | 0.012055 | 16.9861 | 0.28321 | 2.0835 | 1.2446 | -0.00020573 | 0.085585 | 0.011541 | 0.000479 | 0.0021067 | 0.0044094 | 0.00011261 | -0.0048934 | 1.0631 | 0.16482 | 0.01063 | 0.008864 | 0.010833 | 0.0098243
7014678 | 0.005534 | 0.011013 | 16.9327 | 0.27958 | 2.0892 | 1.2502 | 0.00045355 | 0.08465 | 0.011433 | 0.00070424 | 0.0021988 | 0.0049593 | 0.00012152 | -0.0044986 | 1.0409 | 0.15935 | 0.010793 | 0.010498 | 0.011568 | 0.0094401
8022879 | 0.01869 | 0.10371 | 21.125 | 0.30462 | 12.2319 | 1.0475 | 0.00059769 | 0.30852 | 0.085159 | 0.003279 | 0.002003 | 0.059661 | 0.00042799 | 0.00027422 | 1.3652 | 0.79794 | 0.012493 | 0.0061989 | 0.011666 | 0.0071737
8022879 | 0.017864 | 0.097986 | 19.8242 | 0.28555 | 11.4193 | 0.97304 | 0.00033525 | 0.28924 | 0.079166 | 0.002911 | 0.0021233 | 0.049693 | 0.00044897 | 0.00030909 | 1.3118 | 0.75782 | 0.01235 | 0.0068522 | 0.011468 | 0.0068821
3018210 | 0.0015305 | 0.013764 | 5.4508 | 0.1206 | 0.21208 | 0.43869 | -0.00094152 | 0.05399 | 0.0031582 | 0.0003183 | 0.0010442 | 0.020206 | 2.8911e-05 | -0.0060639 | 0.22925 | 0.072262 | 0.0022249 | 0.0018114 | 0.0017361 | 0.001541
3018210 | 0.001801 | 0.012253 | 5.123 | 0.11826 | 0.19933 | 0.43675 | -0.00059997 | 0.052588 | 0.0027988 | 0.00028067 | 0.0012082 | 0.020055 | 5.7391e-05 | -0.0060945 | 0.22128 | 0.071944 | 0.0017842 | 0.0017814 | 0.0013735 | 0.0013746
3014908 | 0.0030009 | 0.049108 | 1.8548 | 0.093522 | 0.67043 | 0.23818 | 0.0052341 | 0.026136 | 0.021032 | 0.00016939 | 0.00070738 | 0.0028078 | 6.3281e-05 | -0.005674 | 0.1514 | 0.028558 | 0.001756 | 0.0015424 | 0.0016088 | 0.0014012
5017696 | 0.013769 | 0.23897 | 36.7971 | 0.37141 | 3.312 | 1.8983 | 0.0037183 | 0.16504 | 0.024291 | 0.0015454 | 0.0077708 | 0.064763 | 0.00017733 | 0.018991 | 1.6862 | 19.9593 | 0.027639 | 0.027183 | 0.025212 | 0.024336
5017696 | 0.012276 | 0.22311 | 34.6143 | 0.34998 | 3.1354 | 1.8278 | 0.0037654 | 0.15761 | 0.024199 | 0.0015822 | 0.0072406 | 0.059859 | 0.00018599 | 0.01659 | 1.5927 | 18.902 | 0.027713 | 0.024661 | 0.025134 | 0.023802
5019710 | 0.020595 | 0.10211 | 6.147 | 0.7778 | 7.4077 | 1.9112 | 0.00050758 | 0.27446 | 0.11466 | 0.0010517 | 0.014458 | 0.02434 | 0.00010431 | 0.0071581 | 3.0408 | 4.1109 | 0.010767 | 0.0073712 | 0.011468 | 0.0065163
5019710 | 0.022666 | 0.1037 | 9.1353 | 0.72061 | 7.9411 | 2.1289 | 0.00075402 | 0.30452 | 0.12747 | 0.0013254 | 0.015649 | 0.024013 | 6.8161e-05 | -0.0008089 | 3.2191 | 4.39 | 0.011293 | 0.0056768 | 0.011423 | 0.0075155
6013988 | 0.007076 | 0.028864 | 24.4573 | 0.26761 | 2.8539 | 0.53756 | 0.0010532 | 0.092389 | 0.020269 | 0.00089787 | 0.0022873 | 0.046571 | 0.00015106 | -0.0021242 | 1.2037 | 0.8002 | 0.0081746 | 0.0064623 | 0.0095439 | 0.0048873
6013988 | 0.0065617 | 0.024042 | 23.7826 | 0.25579 | 2.7617 | 0.52059 | 0.0010701 | 0.08915 | 0.019353 | 0.0010538 | 0.0022294 | 0.04336 | 0.00019163 | -0.0028881 | 1.1454 | 0.76241 | 0.0087718 | 0.0064004 | 0.0086345 | 0.0048874
7017537 | 0.014224 | 0.017745 | 3.5732 | 0.16734 | 1.0368 | 0.5253 | 0.00015018 | 0.052196 | 0.015952 | 0.0061521 | 0.0012585 | 0.010537 | 7.8898e-05 | -0.0046695 | 0.76762 | 0.29051 | 0.0059071 | 0.0064729 | 0.0067257 | 0.0055021
7017537 | 0.012974 | 0.016078 | 3.5822 | 0.16973 | 1.0352 | 0.52086 | 0.00048877 | 0.051082 | 0.015888 | 0.0059527 | 0.0012528 | 0.010791 | 9.8926e-05 | -0.004423 | 0.75043 | 0.29511 | 0.0064658 | 0.0056262 | 0.0063226 | 0.0056031
8015490 | 0.0092471 | 0.077944 | 15.6607 | 0.64484 | 3.0723 | 1.5277 | 0.0017785 | 0.1509 | 0.085238 | 0.0042793 | 0.0046987 | 0.020514 | 0.00067204 | 0.017243 | 1.5498 | 0.090153 | 0.0082251 | 0.012996 | 0.0093667 | 0.0073519
8015490 | 0.0085266 | 0.062808 | 15.961 | 0.66682 | 3.0616 | 1.5359 | 0.0024455 | 0.15004 | 0.085722 | 0.0041785 | 0.0048441 | 0.020446 | 0.00063068 | 0.01664 | 1.5204 | 0.090671 | 0.0096207 | 0.013802 | 0.0095375 | 0.006213
8022828 | 0.011937 | 0.08048 | 31.3906 | 0.38165 | 6.1587 | 1.7299 | 0.0021251 | 0.21777 | 0.10797 | 0.0019566 | 0.0076656 | 0.060226 | 0.00055609 | 0.001553 | 1.8278 | 0.72746 | 0.014427 | 0.011153 | 0.015414 | 0.010945
8022828 | 0.011592 | 0.085539 | 32.0703 | 0.37741 | 6.0036 | 1.7087 | 0.0028197 | 0.21659 | 0.10065 | 0.0022759 | 0.0073728 | 0.061308 | 0.00060054 | 0.0013351 | 1.8427 | 0.73094 | 0.014544 | 0.0096838 | 0.012794 | 0.010329
5019710 | 0.019959 | 0.105 | 6.3289 | 0.77817 | 7.376 | 1.9307 | 0.00097048 | 0.28783 | 0.11847 | 0.0013897 | 0.014349 | 0.024042 | 9.0014e-05 | -0.00013422 | 3.0628 | 4.138 | 0.010516 | 0.0053808 | 0.010103 | 0.0071734
No_sample | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233 | 233
| Mass | Conc. unit | Std1 | Std2 | Std3 | Std4 | Std5 |  | LOD (µg/L) | LOQ (µg/L) |  | LOD Vesna (µg/L)
Co | 58.9332 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.004343 | 0.013161 |  | 0.01832
Ni | 59.9332 | ug/L | 1 | 0.5 | 0.25 | 0.1 | 0.05 |  | 0.026121 | 0.079154 |  | 0.097218
Zn | 65.926 | ug/L | 20 | 10 | 5 | 2 | 1 |  | 0.28966 | 0.87775 |  | 4.3012
Cu | 64.9278 | ug/L | 5 | 2.5 | 1.25 | 0.5 | 0.25 |  | 0.090839 | 0.27527 |  | 0.38529
Sr | 87.9056 | ug/L | 5 | 2.5 | 1.25 | 0.5 | 0.25 |  | 0.080614 | 0.24428 |  | 0.17541
Mo | 97.9055 | ug/L | 10 | 5 | 2.5 | 1 | 0.5 |  | 0.19191 | 0.58154 |  | 0.31149
Sb | 120.904 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.0050962 | 0.015443 |  | 0.030087
Cs | 132.905 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.0036974 | 0.011204 |  | 0.010068
Ba | 137.905 | ug/L | 1 | 0.5 | 0.25 | 0.1 | 0.05 |  | 0.010838 | 0.032843 |  | 0.16255
W | 183.951 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.0062135 | 0.018829 |  | 0.014935
Tl | 204.975 | ug/L | 1 | 0.5 | 0.25 | 0.1 | 0.05 |  | 0.015476 | 0.046896 |  | 0.030419
Pb | 207.977 | ug/L | 1 | 0.5 | 0.25 | 0.1 | 0.05 |  | 0.019132 | 0.057976 |  | 0.027976
U | 238.05 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.0037753 | 0.01144 |  | 0.0076634
Mn | 54.9381 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.0090779 | 0.027509 |  | 0.14609
Se | 77.9173 | ug/L | 5 | 2.5 | 1.25 | 0.5 | 0.25 |  | 0.12208 | 0.36995 |  | 0.51226
AsO | 90.9165 | ug/L | 5 | 2.5 | 1.25 | 0.5 | 0.25 |  | 0.12134 | 0.36768 |  | 0.13538
Cd | 110.904 | ug/L | 0.2 | 0.1 | 0.05 | 0.02 | 0.01 |  | 0.0053262 | 0.01614 |  | 0.026702
Co_STD (µg/L) | Freq | % tot. sample | Ni_STD (µg/L) | Freq | % tot. sample | Zn_STD (µg/L) | Freq | % tot. sample | Cu_STD (µg/L) | Freq | % tot. sample | Sr_STD | Freq | % tot. sample | Mo_STD | Freq | % tot sample | Sb_STD | Freq | % tot sample | Cs_STD | Freq | % tot sample | Ba_STD | Freq | % tot sample | W_STD | Freq | % tot sample | Tl_STD | Freq | % tot sample | Pb_STD | Freq | % tot sample | U_STD | Freq | % tot sample | Mn_STD | Freq | % total sample | Se_STD | Freq | % tot sample | AsO_STD | Freq | % tot sample | Cd_STD | Freq | % tot sample | Cd_1_STD | Freq | Cd_2_STD | Freq
0.01 | 144 | 61.8026 | 0.05 | 160 | 68.6695 | 1 | 16 | 6.867 | 0.25 | 87 | 37.3391 | 0.25 | 11 | 4.721 | 0.5 | 70 | 30.0429 | 0.01 | 232 | 99.5708 | 0.01 | 0 | 0 | 0.05 | 184 | 78.97 | 0.01 | 227 | 97.4249 | 0.05 | 232 | 99.5708 | 0.05 | 209 | 89.6996 | 0.01 | 233 | 100 | 0.01 | 184 | 78.97 | 0.25 | 12 | 5.1502 | 0.25 | 125 | 53.6481 | 0.01 | 100 | 42.9185 | 0.01 | 148 | 0.01 | 105
0.02 | 68 | 29.1845 | 0.1 | 37 | 15.8798 | 2 | 11 | 4.721 | 0.5 | 104 | 44.6352 | 0.5 | 8 | 3.4335 | 1 | 72 | 30.9013 | 0.02 | 0 |  | 0.02 | 4 | 1.7167 | 0.1 | 27 | 11.588 | 0.02 | 2 |  | 0.1 | 0 |  | 0.1 | 21 |  | 0.02 | 0 |  | 0.02 | 30 |  | 0.5 | 42 |  | 0.5 | 39 |  | 0.02 | 87 |  | 0.02 | 42 | 0.02 | 82
0.05 | 18 | 7.7253 | 0.25 | 23 | 9.8712 | 5 | 20 | 8.5837 | 1.25 | 27 | 11.588 | 1.25 | 69 | 29.6137 | 2.5 | 81 | 34.7639 | 0.05 | 0 |  | 0.05 | 42 | 18.0258 | 0.25 | 15 | 6.4378 | 0.05 | 3 |  | 0.25 | 0 |  | 0.25 | 2 |  | 0.05 | 0 |  | 0.05 | 9 |  | 1.25 | 107 |  | 1.25 | 51 |  | 0.05 | 40 |  | 0.05 | 38 | 0.05 | 41
0.1 | 2 | 0.85837 | 0.5 | 0 | 0 | 10 | 55 | 23.6052 | 2.5 | 14 | 6.0086 | 2.5 | 70 | 30.0429 | 5 | 7 | 3.0043 | 0.1 | 0 |  | 0.1 | 65 | 27.897 | 0.5 | 3 | 1.2876 | 0.1 | 0 |  | 0.5 | 0 |  | 0.5 | 0 |  | 0.1 | 0 |  | 0.1 | 5 |  | 2.5 | 60 |  | 2.5 | 9 |  | 0.1 | 5 |  | 0.1 | 4 | 0.1 | 4
0.2 | 0 | 0 | 1 | 0 | 0 | 20 | 84 | 36.0515 | 5 | 0 | 0 | 5 | 60 | 25.7511 | 10 | 3 | 1.2876 | 0.2 | 0 |  | 0.2 | 87 | 37.3391 | 1 | 2 | 0.85837 | 0.2 | 0 |  | 1 | 1 |  | 1 | 0 |  | 0.2 | 0 |  | 0.2 | 2 |  | 5 | 9 |  | 5 | 4 |  | 0.2 | 0 |  | 0.2 | 0 | 0.2 | 0
More | 1 | 0.42918 | More | 13 | 5.5794 | More | 47 | 20.1717 | More | 1 | 0.42918 | More | 15 | 6.4378 | More | 0 |  | More | 1 |  | More | 35 | 15.0215 | More | 2 | 0.85837 | More | 1 |  | More | 0 |  | More | 1 |  | More | 0 |  | More | 3 |  | More | 3 | 1.2876 | More | 5 |  | More | 1 |  | More | 1 | More | 1
= remove standards
Slope_Calibration | Date | Co 59 (ug/L) | Ni 60 (ug/L) | Zn 66 (ug/L) | Cu 65 (ug/L) | Sr 88 (ug/L) | Mo 98 (ug/L) | Sb 121 (ug/L) | Cs 133 (ug/L) | Ba 138 (ug/L) | W 184 (ug/L) | Tl 205 (ug/L) | Pb 208 (ug/L) | U 238 (ug/L) | Mn 55 (ug/L) | Se 78 (ug/L) | AsO 91 (ug/L)
LOD | 2020-01-08 | [11 empty] | 0.023703 | 0.0025013 | 0.0097232 | 0.07167
LOD | 2020-01-08 | [11 empty] | 0.0093437 | 0.0032662 | 0.0060102 | 0.16529
LOD | 2020-01-08 | [11 empty] | 0.022091 | 0.0024751 | 0.0055623 | 0.081088
LOD | 2019-11-20 | [11 empty] | 0.024447 | 0.0050323 | 0.017738 | 0.089943
LOD | 2020-02-27 | [11 empty] | 0.015765 | 0.0041287 | 0.007818 | 0.21402
LOD | 2020-02-27 | [11 empty] | 0.013407 | 0.0033112 | 0.0039524 | 0.11049
Mean (µg/L) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0.018126 | 0.0034525 | 0.0084673 | 0.12208 | #DIV/0!